## Table S1
|  | Data for quality control (QC) samples | [8 empty] | Pooled values within an instrument for all QC |  |  | Pooled values for all instruments and QC
Analyte | Instrument | QC material | QC lot | Number of values | Min value, SI units | Max value, SI units | Mean, SI units | CV % | CV %, lower 95% confidence limit | CV %, upper 95% confidence limit | Pooled CV% | Pooled CV %, lower 95% confidence limit | Pooled CV %, upper 95% confidence limit | Pooled CV% | Pooled CV %, lower 95% confidence limit | Pooled CV %, upper 95% confidence limit
ADENOSINE DEAMINASE FLUID, U/L | C6000-B:C2 | [12 empty] | 4.2331 | 4.0214 | 4.4447
|  | ADA 2 | 201220 | 132 |  |  | 142.541 | 4.45 | 4.2275 | 4.6725
|  | ADA 2 | ADC203210 | 46 |  |  | 140.348 | 4.632 | 4.4004 | 4.8636
|  | ADA 1 | ADC103210 | 117 |  |  | 31.75 | 4.901 | 4.6559 | 5.146
|  | ADA 2 | 202210 | 102 |  |  | 155.545 | 3.846 | 3.6537 | 4.0383
|  | ADA 0 | 01220 | 131 |  |  | 11.458 | 5.337 | 5.0701 | 5.6039
|  | ADA 0 | ADCL02210 | 106 |  |  | 11.259 | 3.044 | 2.8918 | 3.1962
|  | ADA 1 | 101220 | 96 |  |  | 30.457 | 4.017 | 3.8161 | 4.2179
|  | ADA 1 | 102210 | 67 |  |  | 32.54 | 2.964 | 2.8158 | 3.1122
|  | ADA 0 | L01210 | 49 |  |  | 10.698 | 2.775 | 2.6362 | 2.9137
ADRENOCORTICOTROPIC HORMONE, U/L | C6000-A:E1-1 | [12 empty] | 4.9617 | 4.7136 | 5.2098
|  | PC MM1 | 523425 | 40 |  |  | 8.821 | 5.165 | 4.9067 | 5.4233
|  | PC MM2 | 523426 | 39 |  |  | 194.382 | 4.744 | 4.5068 | 4.9812
ALANINE TRANSAMINASE, U/L | C6000-A:C1 | [12 empty] | 3.8169 | 3.6261 | 4.0078
|  | PCCC 1 | 525027 | 329 |  |  | 45.889 | 2.932 | 2.7854 | 3.0786
|  | PCCC 2 | 535719 | 93 |  |  | 122.088 | 2.783 | 2.6439 | 2.9221
|  | s-PCCC 2 | 461597 | 39 |  |  | 121.946 | 2.104 | 1.9988 | 2.2092
|  | PCCC 2 | 519198 | 369 |  |  | 116.679 | 2.422 | 2.3009 | 2.5431
|  | s-PCCC 1 | 525027 | 47 |  |  | 47.455 | 2.757 | 2.6191 | 2.8948
|  | PCCC 1 | 525205 | 147 |  |  | 47.612 | 1.929 | 1.8325 | 2.0255
ALANINE TRANSAMINASE, U/L | C6000-B:C2 | [12 empty] | 3.8169 | 3.6261 | 4.0078
|  | PCCC 1-B | 525205 | 158 |  |  | 47.963 | 2.775 | 2.6362 | 2.9137
|  | s-PCCC 2-B | 461597 | 41 |  |  | 118.668 | 1.212 | 1.1514 | 1.2726
|  | PCCC 2-B | 535719 | 98 |  |  | 122.306 | 10.557 | 10.0291 | 11.0848
|  | PCCC 2-B | 519198 | 365 |  |  | 116.165 | 3.943 | 3.7458 | 4.1402
|  | s-PCCC 1-B | 525027 | 46 |  |  | 46.1 | 1.53 | 1.4535 | 1.6065
|  | PCCC 1-B | 525027 | 311 |  |  | 45.36 | 4.014 | 3.8133 | 4.2147
ALBUMIN, U/L | C6000-A:C1 | [12 empty] | 2.8601 | 2.7171 | 3.0031
|  | PCCC 1 | 525205 | 141 |  |  | 32.523 | 3.197 | 3.0372 | 3.3569
|  | PCCC 2 | 519198 | 371 |  |  | 49.789 | 1.831 | 1.7394 | 1.9225
|  | PCCC 2 | 535719 | 88 |  |  | 48.601 | 2.165 | 2.0568 | 2.2732
|  | PCCC 2 | 461597 | 73 |  |  | 51.996 | 2.014 | 1.9133 | 2.1147
|  | PCCC 1 | 525027 | 407 |  |  | 31.602 | 2.886 | 2.7417 | 3.0303
ALBUMIN, U/L | C6000-B:C2 | [12 empty] | 2.8601 | 2.7171 | 3.0031
|  | s-PCCC 2-B | 519198 | 584 |  |  | 49.95 | 2.779 | 2.6401 | 2.9179
|  | s-PCCC 1-B | 525205 | 227 |  |  | 31.924 | 4.315 | 4.0993 | 4.5308
|  | s-PCCC 2-B | 535719 | 146 |  |  | 47.631 | 2.319 | 2.2031 | 2.435
|  | PCCC 2-B | 461597 | 68 |  |  | 51.828 | 1.861 | 1.7679 | 1.9541
|  | s-PCCC 1-B | 525027 | 488 |  |  | 31.792 | 3.183 | 3.0238 | 3.3422
|  | PCCC 1-B | 525027 | 78 |  |  | 30.965 | 2.133 | 2.0263 | 2.2397
ALBUMIN CSF, U/L | C6000-A:C1 | [12 empty] | 3.8053 | 3.615 | 3.9955
|  | LIQ SFC2 | 56502 | 110 |  |  | 342.374 | 3.435 | 3.2633 | 3.6067
|  | LIQ SFC2 | 56492 | 48 |  |  | 350.658 | 1.765 | 1.6767 | 1.8533
|  | PPPUC TPC3 | 646927 | 121 |  |  | 121.158 | 3.996 | 3.7962 | 4.1958
|  | LIQ SFC2 | 56522 | 128 |  |  | 317.951 | 3.656 | 3.4732 | 3.8388
|  | PPPUC TPC3 | 637091 | 120 |  |  | 121.775 | 4.929 | 4.6825 | 5.1755
|  | PPPUC TPC3 | 531181 | 37 |  |  | 119.67 | 2.911 | 2.7654 | 3.0566
|  | LIQ SFC2 | 56532 | 39 |  |  | 306.392 | 2.941 | 2.794 | 3.088
|  | PPPUC TPC3 | 602710 | 157 |  |  | 127.754 | 4.332 | 4.1154 | 4.5486
|  | LIQ SFC2 | 56512 | 100 |  |  | 353.524 | 2.885 | 2.7407 | 3.0293
ALBUMIN URINE, U/L | C6000-A:C1 | [12 empty] | 10.5758 | 10.047 | 11.1046
|  | PN PUC | 602712 | 123 |  |  | 31.905 | 3.774 | 3.5853 | 3.9627
|  | PP PUC | 64692700 | 84 |  |  | 114.132 | 2.439 | 2.3171 | 2.561
|  | PN PUC | 649811 | 183 |  |  | 31.864 | 18.891 | 17.9464 | 19.8356
|  | Ran-Liq 2 | 1237UC | 75 |  |  | 25.111 | 22.904 | 21.7588 | 24.0492
|  | PN PUC | 531179 | 82 |  |  | 27.938 | 3.603 | 3.4228 | 3.7832
|  | PP PUC | 531181 | 40 |  |  | 114.28 | 1.82 | 1.729 | 1.911
|  | Ran-Liq 2 | 1209UC | 68 |  |  | 23.709 | 9.198 | 8.7381 | 9.6579
|  | Ran-Liq 3 | 1238UC | 79 |  |  | 63.152 | 6.381 | 6.062 | 6.7001
|  | Ran-Liq 3 | 1210UC | 68 |  |  | 63.488 | 3.215 | 3.0543 | 3.3758
|  | PP PUC | 637091 | 121 |  |  | 119.967 | 3.508 | 3.3326 | 3.6834
|  | PP PUC | 602710 | 161 |  |  | 122.05 | 3.37 | 3.2015 | 3.5385
ALDO, U/L | LIAISON XL | [12 empty] | 8.3729 | 7.9542 | 8.7915
|  | ALDO L1 | 227012 | 51 |  |  | 189.294 | 10.004 | 9.5038 | 10.5042
|  | ALDO L2 | 228012 | 48 |  |  | 740.333 | 6.182 | 5.8729 | 6.4911
ALKALINE PHOSPHATASE, U/L | C6000-A:C1 | [12 empty] | 4.8814 | 4.6374 | 5.1255
|  | s-PCCC 2 | 519198 | 619 |  |  | 239.817 | 4.231 | 4.0194 | 4.4425
|  | PCCC 2 | 461597 | 69 |  |  | 231.623 | 2.933 | 2.7864 | 3.0796
|  | s-PCCC 1 | 525205 | 268 |  |  | 100.791 | 5.523 | 5.2468 | 5.7992
|  | PCCC 1 | 525027 | 92 |  |  | 90.413 | 4.647 | 4.4146 | 4.8794
|  | s-PCCC 1 | 525027 | 520 |  |  | 91.083 | 5.308 | 5.0426 | 5.5734
|  | s-PCCC 2 | 535719 | 160 |  |  | 267.206 | 5.471 | 5.1974 | 5.7446
ALKALINE PHOSPHATASE, U/L | C6000-B:C2 | [12 empty] | 4.8814 | 4.6374 | 5.1255
|  | s-PCCC 2-B | 519198 | 588 |  |  | 237.776 | 4.367 | 4.1486 | 4.5854
|  | PCCC 2-B | 461597 | 68 |  |  | 230.235 | 2.774 | 2.6353 | 2.9127
|  | s-PCCC 1-B | 525205 | 254 |  |  | 97.413 | 6.573 | 6.2443 | 6.9017
|  | s-PCCC 1-B | 525027 | 484 |  |  | 90.426 | 4.822 | 4.5809 | 5.0631
|  | PCCC 1-B | 525027 | 80 |  |  | 88.175 | 5.11 | 4.8545 | 5.3655
|  | s-PCCC 2-B | 535719 | 162 |  |  | 258.068 | 4.17 | 3.9615 | 4.3785
ALPHA 1-ANTITRYPSIN, U/L | C6000-B:C2 | [12 empty] | 2.6384 | 2.5065 | 2.7703
|  | PCCC 1-B | 525027 | 231 |  |  | 0.802 | 3.023 | 2.8718 | 3.1742
|  | PCCC 1-B | 525205 | 139 |  |  | 0.904 | 2.436 | 2.3142 | 2.5578
|  | s-PCCC 2-B | 461597 | 33 |  |  | 1.277 | 1.568 | 1.4896 | 1.6464
|  | s-PCCC 1-B | 525027 | 36 |  |  | 0.802 | 1.738 | 1.6511 | 1.8249
|  | PCCC 2-B | 519198 | 277 |  |  | 1.289 | 2.863 | 2.7199 | 3.0061
|  | PCCC 2-B | 535719 | 85 |  |  | 1.343 | 1.409 | 1.3385 | 1.4794
ALPHA-FETO PROTEIN, U/L | C6000-A:E1-1 | [12 empty] | 5.0977 | 4.8428 | 5.3525
|  | RIAP L3 | 2147 | 158 |  |  | 231.45 | 6.611 | 6.2804 | 6.9416
|  | RIAP L3 | 2181EC | 145 |  |  | 205.785 | 4.099 | 3.8941 | 4.304
|  | RIAP L1 | 2179EC | 139 |  |  | 10.187 | 5.934 | 5.6373 | 6.2307
|  | RIAP L1 | 2105EC | 311 |  |  | 11.715 | 4.779 | 4.54 | 5.0179
|  | RIAP L3 | 2110EC | 149 |  |  | 231.58 | 3.757 | 3.5692 | 3.9449
AMMONIA, U/L | C6000-B:C2 | [12 empty] | 3.3831 | 3.214 | 3.5523
|  | AMM-P | 623128 | 81 |  |  | 233.749 | 2.111 | 2.0055 | 2.2165
|  | AMM-N | 644004 | 167 |  |  | 63.522 | 3.62 | 3.439 | 3.801
|  | AMM-N | 587880 | 232 |  |  | 63.501 | 4.143 | 3.9358 | 4.3502
|  | AMM-P | 587881 | 227 |  |  | 223.73 | 2.859 | 2.7161 | 3.0019
|  | AMM-P | 644006 | 87 |  |  | 227.843 | 2.825 | 2.6837 | 2.9663
ANTI-THYROGLOBULIN AB, U/L | C6000-A:E2-1 | [12 empty] | 6.6651 | 6.3318 | 6.9983
|  | PC THYRO1 | 580444 | 228 |  |  | 112.67 | 7.186 | 6.8267 | 7.5453
|  | PC THYRO2 | 580446 | 231 |  |  | 221.702 | 7.653 | 7.2703 | 8.0357
|  | PC THYRO2 | 612140 | 110 |  |  | 217.754 | 5.703 | 5.4178 | 5.9882
|  | PC THYRO1 | 612139 | 87 |  |  | 101.275 | 3.899 | 3.7041 | 4.094
|  | PC THYRO1 | 669238 | 49 |  |  | 62.587 | 4.679 | 4.4451 | 4.913
ANTI-THYROID PEROXIDASE AB, U/L | C6000-A:E1-1 | [12 empty] | 8.848 | 8.4056 | 9.2904
|  | PC THYRO2 | 580446 | 42 |  |  | 109.415 | 7.653 | 7.2703 | 8.0357
|  | PC THYRO1 | 580444 | 43 |  |  | 47.726 | 9.876 | 9.3822 | 10.3698
APOPROTEIN A, U/L | C6000-B:C2 | [12 empty] | 4.5237 | 4.2975 | 4.7499
|  | PCCC 1-B | 525027 | 98 |  |  | 1.119 | 4.366 | 4.1477 | 4.5843
|  | PCCC 2-B | 519198 | 116 |  |  | 1.747 | 4.51 | 4.2845 | 4.7355
|  | PCCC 1-B | 525205 | 38 |  |  | 1.153 | 4.952 | 4.7044 | 5.1996
APOPROTEIN B, U/L | C6000-B:C2 | [12 empty] | 5.0186 | 4.7676 | 5.2695
|  | PCCC 1-B | 525205 | 39 |  |  | 0.475 | 5.424 | 5.1528 | 5.6952
|  | PCCC 2-B | 519198 | 110 |  |  | 0.764 | 4.809 | 4.5686 | 5.0495
|  | PCCC 1-B | 525027 | 93 |  |  | 0.471 | 5.087 | 4.8326 | 5.3414
ASPARTATE TRANSAMINASE, U/L | C6000-A:C1 | [12 empty] | 3.6446 | 3.4624 | 3.8268
|  | PCCC 1 | 525205 | 141 |  |  | 44.284 | 1.641 | 1.559 | 1.723
|  | PCCC 2 | 535719 | 91 |  |  | 137.558 | 1.763 | 1.6748 | 1.8512
|  | s-PCCC 1 | 525027 | 43 |  |  | 46 | 2.012 | 1.9114 | 2.1126
|  | PCCC 2 | 519198 | 344 |  |  | 143.651 | 2.16 | 2.052 | 2.268
|  | s-PCCC 2 | 461597 | 37 |  |  | 148.027 | 1.37 | 1.3015 | 1.4385
|  | PCCC 1 | 525027 | 307 |  |  | 45.498 | 2.42 | 2.299 | 2.541
ASPARTATE TRANSAMINASE, U/L | C6000-B:C2 | [12 empty] | 3.6446 | 3.4624 | 3.8268
|  | s-PCCC 2-B | 461597 | 33 |  |  | 149.979 | 2.409 | 2.2885 | 2.5295
|  | PCCC 1-B | 525205 | 139 |  |  | 44.905 | 2.804 | 2.6638 | 2.9442
|  | PCCC 2-B | 535719 | 88 |  |  | 139.222 | 11.021 | 10.47 | 11.5721
|  | PCCC 2-B | 519198 | 302 |  |  | 146.775 | 3.509 | 3.3335 | 3.6845
|  | s-PCCC 1-B | 525027 | 38 |  |  | 46.361 | 2.259 | 2.146 | 2.372
|  | PCCC 1-B | 525027 | 264 |  |  | 46.514 | 3.943 | 3.7458 | 4.1402
BETA-HCG, U/L | C6000-A:E1-1 | [12 empty] | 5.7593 | 5.4713 | 6.0473
|  | RIAP L3 | 2110EC | 151 |  |  | 300.732 | 4.254 | 4.0413 | 4.4667
|  | RIAP L3 | 2181EC | 151 |  |  | 252.719 | 5.082 | 4.8279 | 5.3361
|  | RIAP L1 | 2179EC | 147 |  |  | 11.284 | 5.528 | 5.2516 | 5.8044
|  | RIAP L1 | 2105EC | 324 |  |  | 11.02 | 4.668 | 4.4346 | 4.9014
|  | RIAP L3 | 2147 | 165 |  |  | 295.538 | 4.178 | 3.9691 | 4.3869
BETA-HCG, U/L | C6000-B:E3-2 | [12 empty] | 5.7593 | 5.4713 | 6.0473
|  | RIAP L3-B | 2147EC | 148 |  |  | 305.552 | 6.47 | 6.1465 | 6.7935
|  | RIAP L3-B | 2181EC | 154 |  |  | 258.142 | 6.555 | 6.2272 | 6.8827
|  | RIAP L3-B | 2110EC | 144 |  |  | 294.176 | 4.566 | 4.3377 | 4.7943
|  | RIAP L1-B | 2105EC | 292 |  |  | 11.134 | 7.94 | 7.543 | 8.337
|  | RIAP L1-B | 2179EC | 141 |  |  | 11.468 | 5.935 | 5.6382 | 6.2317
BICARBONATE, U/L | C6000-B:C2 | [12 empty] | 4.4566 | 4.2338 | 4.6795
|  | AMM-P | 644006 | 94 |  |  | 29.154 | 4.887 | 4.6426 | 5.1313
|  | AMM-P | 623128 | 65 |  |  | 29.68 | 3.804 | 3.6138 | 3.9942
|  | AMM-N | 587880 | 219 |  |  | 18.775 | 4.959 | 4.711 | 5.2069
|  | AMM-P | 587881 | 214 |  |  | 32.164 | 3.387 | 3.2176 | 3.5564
|  | AMM-N | 644004 | 160 |  |  | 17.956 | 4.943 | 4.6958 | 5.1902
C-REACTIVE PROTEIN, U/L | C6000-A:C1 | [12 empty] | 4.7053 | 4.47 | 4.9405
|  | s-PCCC 1 | 525027 | 48 |  |  | 8.805 | 2.74 | 2.603 | 2.877
|  | PCCC 2 | 519198 | 366 |  |  | 50.944 | 3.259 | 3.096 | 3.4219
|  | PCCC 1 | 525027 | 321 |  |  | 8.809 | 3.247 | 3.0846 | 3.4093
|  | s-PCCC 2 | 461597 | 43 |  |  | 52.091 | 4.701 | 4.4659 | 4.936
|  | PCCC 2 | 535719 | 98 |  |  | 51.359 | 2.707 | 2.5716 | 2.8424
|  | PCCC 1 | 525205 | 156 |  |  | 5.663 | 2.276 | 2.1622 | 2.3898
C-REACTIVE PROTEIN, U/L | C6000-B:C2 | [12 empty] | 4.7053 | 4.47 | 4.9405
|  | PCCC 2-B | 519198 | 316 |  |  | 51.332 | 3.251 | 3.0884 | 3.4135
|  | s-PCCC 1-B | 525027 | 41 |  |  | 8.731 | 2.074 | 1.9703 | 2.1777
|  | PCCC 1-B | 525027 | 261 |  |  | 8.765 | 4.046 | 3.8437 | 4.2483
|  | PCCC 1-B | 525205 | 139 |  |  | 5.53 | 9.803 | 9.3129 | 10.2932
|  | s-PCCC 2-B | 461597 | 37 |  |  | 54.266 | 2.549 | 2.4215 | 2.6764
|  | PCCC 2-B | 535719 | 85 |  |  | 50.955 | 11.31 | 10.7445 | 11.8755
C-peptide, U/L | C6000-A:E2-2 | [12 empty] | 4.7123 | 4.4767 | 4.9479
|  | RIAP L2 | 2107EC | 357 |  |  | 6.266 | 3.51 | 3.3345 | 3.6855
|  | RIAP L1 | 2105EC | 382 |  |  | 2.103 | 4.315 | 4.0993 | 4.5308
|  | RIAP L1 | 2179EC | 185 |  |  | 1.776 | 7.279 | 6.915 | 7.6429
|  | RIAP L2 | 2124EC | 105 |  |  | 6.684 | 4.689 | 4.4546 | 4.9234
|  | RIAP L2 | 2228EC | 89 |  |  | 6.068 | 3.692 | 3.5074 | 3.8766
CA 125, U/L | C6000-A:E2-1 | [12 empty] | 4.2448 | 4.0326 | 4.4571
|  | RIAP L1 | 2179EC | 121 |  |  | 20.111 | 3.863 | 3.6698 | 4.0562
|  | RIAP L3 | 2147 | 155 |  |  | 190.874 | 5.527 | 5.2507 | 5.8034
|  | RIAP L1 | 2105EC | 302 |  |  | 18.505 | 4.625 | 4.3937 | 4.8563
|  | RIAP L3 | 2110EC | 139 |  |  | 187.761 | 3.253 | 3.0903 | 3.4156
|  | RIAP L3 | 2181EC | 133 |  |  | 195.694 | 2.51 | 2.3845 | 2.6355
CA 19-9, U/L | C6000-A:E2-1 | [12 empty] | 5.5818 | 5.3027 | 5.8609
|  | RIAP L3 | 2110EC | 146 |  |  | 168.101 | 3.852 | 3.6594 | 4.0446
|  | RIAP L1 | 2179EC | 124 |  |  | 19.958 | 5.767 | 5.4787 | 6.0554
|  | RIAP L1 | 2105EC | 327 |  |  | 23.598 | 6.566 | 6.2377 | 6.8943
|  | RIAP L3 | 2181EC | 138 |  |  | 166.654 | 4.298 | 4.0831 | 4.5129
|  | RIAP L3 | 2147 | 164 |  |  | 162.957 | 5.553 | 5.2754 | 5.8307
CALCIUM, U/L | C6000-A:C1 | [12 empty] | 2.9214 | 2.7753 | 3.0675
|  | PCCC 2 | 519198 | 378 |  |  | 3.395 | 1.949 | 1.8516 | 2.0465
|  | PCCC 2 | 535719 | 94 |  |  | 3.458 | 1.453 | 1.3803 | 1.5256
|  | PCCC 2 | 461597 | 70 |  |  | 3.391 | 0.904 | 0.8588 | 0.9492
|  | PCCC 1 | 525205 | 148 |  |  | 2.146 | 2.009 | 1.9085 | 2.1094
|  | PCCC 1 | 525027 | 409 |  |  | 2.237 | 1.801 | 1.7109 | 1.891
CALCIUM, U/L | C6000-B:C2 | [12 empty] | 2.9214 | 2.7753 | 3.0675
|  | PCCC 2-B | 519198 | 371 |  |  | 3.35 | 5.748 | 5.4606 | 6.0354
|  | PCCC 2-B | 535719 | 96 |  |  | 3.441 | 1.037 | 0.98515 | 1.0889
|  | PCCC 2-B | 461597 | 74 |  |  | 3.397 | 1.752 | 1.6644 | 1.8396
|  | PCCC 1-B | 525205 | 156 |  |  | 2.117 | 2.176 | 2.0672 | 2.2848
|  | PCCC 1-B | 525027 | 393 |  |  | 2.22 | 2.089 | 1.9846 | 2.1934
CALCIUM URINE, U/L | C6000-B:C2 | [12 empty] | 2.4815 | 2.3575 | 2.6056
|  | Ran-Liq 3 | 1238UC | 127 |  |  | 2.859 | 2.866 | 2.7227 | 3.0093
|  | Ran-Liq 2 | 1209UC | 224 |  |  | 1.979 | 2.219 | 2.108 | 2.33
|  | Ran-Liq 3 | 1210UC | 224 |  |  | 2.897 | 2.505 | 2.3798 | 2.6303
|  | Bio-liq1 | 88161 | 122 |  |  | 1.826 | 2.447 | 2.3247 | 2.5694
|  | Bio-liq 2 | 88162 | 112 |  |  | 2.71 | 1.74 | 1.653 | 1.827
|  | Ran-Liq 2 | 1237UC | 116 |  |  | 1.925 | 3.045 | 2.8927 | 3.1972
CARCINOEMBRYONIC AG, U/L | C6000-A:E1-2 | [12 empty] | 5.6745 | 5.3908 | 5.9583
|  | RIAP L1 | 2105EC | 314 |  |  | 4.379 | 3.964 | 3.7658 | 4.1622
|  | RIAP L1 | 2179EC | 120 |  |  | 4.204 | 6.127 | 5.8206 | 6.4333
|  | RIAP L3 | 2181EC | 134 |  |  | 50.868 | 8.219 | 7.808 | 8.6299
|  | RIAP L3 | 2147 | 151 |  |  | 52.876 | 6.699 | 6.364 | 7.0339
|  | RIAP L3 | 2110EC | 143 |  |  | 53.743 | 4.191 | 3.9815 | 4.4005
CHLORIDE, U/L | C6000-A:ISE1 | [12 empty] | 3.3173 | 3.1515 | 3.4832
|  | PCCC 1 | 525027 | 106 |  |  | 87.111 | 1.611 | 1.5305 | 1.6916
|  | PCCC 2 | 461597 | 88 |  |  | 103.232 | 1.747 | 1.6597 | 1.8343
|  | s-PCCC 1 | 525205 | 342 |  |  | 84.446 | 2.174 | 2.0653 | 2.2827
|  | s-PCCC 2 | 519198 | 712 |  |  | 105.235 | 1.747 | 1.6597 | 1.8343
|  | s-PCCC 1 | 525027 | 595 |  |  | 87.555 | 2.184 | 2.0748 | 2.2932
|  | s-PCCC 2 | 535719 | 193 |  |  | 105.448 | 1.658 | 1.5751 | 1.7409
CHLORIDE, U/L | C6000-B:ISE1 | [12 empty] | 3.3173 | 3.1515 | 3.4832
|  | PCCC 1-B | 525027 | 368 |  |  | 88.191 | 1.762 | 1.6739 | 1.8501
|  | PCCC 2-B | 519198 | 278 |  |  | 105.84 | 1.499 | 1.4241 | 1.5739
|  | PCCC 2-B | 461597 | 92 |  |  | 103.572 | 1.328 | 1.2616 | 1.3944
|  | s-PCCC 2-B | 535719 | 188 |  |  | 106.224 | 3.356 | 3.1882 | 3.5238
|  | s-PCCC 1-B | 525027 | 565 |  |  | 88.418 | 1.773 | 1.6844 | 1.8617
|  | s-PCCC 1-B | 525205 | 303 |  |  | 85.98 | 10.585 | 10.0557 | 11.1143
|  | s-PCCC 2-B | 519198 | 696 |  |  | 106.061 | 1.837 | 1.7451 | 1.9288
CHLORIDE URINE, U/L | C6000-B:ISE1 | [12 empty] | 3.2449 | 3.0826 | 3.4071
|  | Bio-liq1 | 88161 | 165 |  |  | 93.919 | 2.229 | 2.1176 | 2.3405
|  | Bio-liq 2 | 88162 | 157 |  |  | 194.188 | 2.051 | 1.9485 | 2.1536
|  | Ran-Liq 3 | 1238UC | 154 |  |  | 180.447 | 2.614 | 2.4833 | 2.7447
|  | Ran-Liq 3 | 1210UC | 252 |  |  | 168.146 | 3.876 | 3.6822 | 4.0698
|  | Ran-Liq 2 | 1237UC | 138 |  |  | 79.582 | 2.111 | 2.0055 | 2.2165
|  | Ran-Liq 2 | 1209UC | 246 |  |  | 68.199 | 4.402 | 4.1819 | 4.6221
COMPLEMENT C3, U/L | C6000-B:C2 | [12 empty] | 3.7153 | 3.5295 | 3.9011
|  | PCCC 2-B | 519198 | 248 |  |  | 1.588 | 3.932 | 3.7354 | 4.1286
|  | s-PCCC 1-B | 525027 | 30 |  |  | 0.961 | 1.612 | 1.5314 | 1.6926
|  | PCCC 1-B | 525027 | 212 |  |  | 0.974 | 3.035 | 2.8832 | 3.1867
|  | PCCC 1-B | 525205 | 109 |  |  | 0.938 | 4.16 | 3.952 | 4.368
|  | PCCC 2-B | 535719 | 65 |  |  | 1.507 | 4.66 | 4.427 | 4.893
COMPLEMENT C4, U/L | C6000-B:C2 | [12 empty] | 2.969 | 2.8205 | 3.1174
|  | PCCC 1-B | 525205 | 109 |  |  | 0.139 | 2.196 | 2.0862 | 2.3058
|  | PCCC 1-B | 525027 | 195 |  |  | 0.124 | 3.519 | 3.343 | 3.6949
|  | PCCC 2-B | 519198 | 231 |  |  | 0.218 | 3.09 | 2.9355 | 3.2445
|  | PCCC 2-B | 535719 | 66 |  |  | 0.223 | 1.484 | 1.4098 | 1.5582
CONJUGATED BILIRUBIN, U/L | C6000-A:C1 | [12 empty] | 4.5514 | 4.3239 | 4.779
|  | PCCC 2 | 461597 | 69 |  |  | 41.146 | 4.491 | 4.2664 | 4.7155
|  | PCCC 1 | 525027 | 393 |  |  | 14.943 | 6.381 | 6.062 | 6.7001
|  | MAS BILIRUBIN 3 | BC24033A | 386 |  |  | 172.779 | 4.039 | 3.8371 | 4.2409
|  | PCCC 1 | 525205 | 151 |  |  | 16.193 | 3.551 | 3.3735 | 3.7285
|  | PCCC 2 | 519198 | 354 |  |  | 41.764 | 4.276 | 4.0622 | 4.4898
|  | PCCC 2 | 535719 | 98 |  |  | 44.05 | 2.994 | 2.8443 | 3.1437
|  | MAS BILIRUBIN 3 | BC23063A | 66 |  |  | 167.064 | 4.686 | 4.4517 | 4.9203
CONJUGATED BILIRUBIN, U/L | C6000-B:C2 | [12 empty] | 4.5514 | 4.3239 | 4.779
|  | PCCC 1-B | 525027 | 354 |  |  | 15.474 | 4.739 | 4.502 | 4.976
|  | PCCC 2-B | 535719 | 93 |  |  | 43.499 | 3.587 | 3.4076 | 3.7664
|  | PCCC 2-B | 519198 | 328 |  |  | 42.629 | 4.937 | 4.6902 | 5.1839
|  | PCCC 1-B | 525205 | 148 |  |  | 16.161 | 3.781 | 3.592 | 3.9701
|  | MAS BILIRUBIN-B | BC24033A | 338 |  |  | 174.774 | 4.088 | 3.8836 | 4.2924
|  | PCCC 2-B | 461597 | 67 |  |  | 41.71 | 2.253 | 2.1404 | 2.3657
|  | MAS BILIRUBIN-B | BC23063A | 61 |  |  | 170.185 | 2.896 | 2.7512 | 3.0408
CORTISOL II, U/L | C6000-A:E1-2 | [12 empty] | 4.6383 | 4.4064 | 4.8702
|  | RIAP L1 | 2105EC | 300 |  |  | 197.438 | 4.144 | 3.9368 | 4.3512
|  | RIAP L2 | 2228EC | 63 |  |  | 694.73 | 6.147 | 5.8396 | 6.4544
|  | RIAP L1 | 2179EC | 117 |  |  | 198.93 | 5.838 | 5.5461 | 6.1299
|  | RIAP L2 | 2124EC | 74 |  |  | 641.069 | 5.738 | 5.4511 | 6.0249
|  | RIAP L2 | 2107EC | 281 |  |  | 643.459 | 3.79 | 3.6005 | 3.9795
CREATINE KINASE Gen 2, U/L | C6000-A:C1 | [12 empty] | 1.8516 | 1.759 | 1.9442
|  | s-PCCC 2 | 461597 | 36 |  |  | 259.417 | 1.368 | 1.2996 | 1.4364
|  | PCCC 1 | 525027 | 289 |  |  | 149.512 | 2.159 | 2.0511 | 2.267
|  | s-PCCC 1 | 525027 | 44 |  |  | 149.636 | 1.117 | 1.0612 | 1.1728
|  | PCCC 1 | 525205 | 134 |  |  | 148.224 | 1.401 | 1.331 | 1.471
|  | PCCC 2 | 535719 | 82 |  |  | 261.11 | 1.443 | 1.3708 | 1.5151
|  | PCCC 2 | 519198 | 329 |  |  | 256.502 | 1.93 | 1.8335 | 2.0265
CREATININE, U/L | C6000-A:C1 | [12 empty] | 2.14 | 2.033 | 2.247
|  | PCCC 2 | 461597 | 136 |  |  | 361.184 | 1.648 | 1.5656 | 1.7304
|  | PCCC 1 | 525205 | 417 |  |  | 89.326 | 2.232 | 2.1204 | 2.3436
|  | PCCC 1 | 525027 | 979 |  |  | 89.396 | 2.292 | 2.1774 | 2.4066
|  | PCCC 2 | 535719 | 259 |  |  | 330.95 | 2.131 | 2.0244 | 2.2376
|  | PCCC 2 | 519198 | 982 |  |  | 352.287 | 1.852 | 1.7594 | 1.9446
CREATININE, U/L | C6000-B:C2 | [12 empty] | 2.14 | 2.033 | 2.247
|  | PCCC 1-B | 525027 | 762 |  |  | 88.979 | 2.387 | 2.2677 | 2.5063
|  | PCCC 2-B | 519198 | 759 |  |  | 347.335 | 2.265 | 2.1517 | 2.3782
|  | PCCC 2-B | 461597 | 121 |  |  | 351.909 | 2.525 | 2.3988 | 2.6513
|  | PCCC 1-B | 525205 | 359 |  |  | 87.518 | 1.851 | 1.7585 | 1.9436
|  | PCCC 2-B | 535719 | 231 |  |  | 322.996 | 1.541 | 1.464 | 1.618
CREATININE URINE, U/L | C6000-A:C1 | [12 empty] | 2.7193 | 2.5833 | 2.8552
|  | Ran-Liq 3 | 1210UC | 119 |  |  | 14.508 | 3.757 | 3.5692 | 3.9449
|  | Ran-Liq 2 | 1237UC | 114 |  |  | 6.48 | 2.953 | 2.8053 | 3.1006
|  | PP PUC | 64692700 | 186 |  |  | 4.558 | 2.128 | 2.0216 | 2.2344
|  | PN PUC | 649811 | 401 |  |  | 8.843 | 2.072 | 1.9684 | 2.1756
|  | Ran-Liq 2 | 1209UC | 124 |  |  | 6.357 | 3.352 | 3.1844 | 3.5196
|  | PN PUC | 602712 | 295 |  |  | 8.525 | 2.746 | 2.6087 | 2.8833
|  | PN PUC | 531179 | 174 |  |  | 8.619 | 3.803 | 3.6128 | 3.9931
|  | PN PUC | 634291 | 37 |  |  | 8.92 | 1.365 | 1.2968 | 1.4332
|  | PP PUC | 602710 | 385 |  |  | 4.531 | 2.867 | 2.7237 | 3.0103
|  | PP PUC | 531181 | 80 |  |  | 4.451 | 3.702 | 3.5169 | 3.8871
|  | Ran-Liq 3 | 1238UC | 112 |  |  | 14.759 | 2.615 | 2.4842 | 2.7458
|  | PP PUC | 637091 | 239 |  |  | 4.522 | 2.125 | 2.0187 | 2.2313
CREATININE URINE, U/L | C6000-B:C2 | [12 empty] | 2.7193 | 2.5833 | 2.8552
|  | Ran-Liq 3 | 1238UC | 33 |  |  | 15.144 | 1.196 | 1.1362 | 1.2558
|  | Ran-Liq 2 | 1237UC | 145 |  |  | 6.536 | 2.126 | 2.0197 | 2.2323
|  | Ran-Liq 2 | 1209UC | 299 |  |  | 6.431 | 2.609 | 2.4785 | 2.7395
DEHYDROEPIANDROSTERONE SULPHAT, U/L | C6000-A:E2-2 | [12 empty] | 5.5647 | 5.2864 | 5.8429
|  | RIAP L2 | 2107EC | 277 |  |  | 15.98 | 4.841 | 4.599 | 5.0831
|  | RIAP L1 | 2179EC | 117 |  |  | 2.635 | 6.794 | 6.4543 | 7.1337
|  | RIAP L2 | 2124EC | 80 |  |  | 16.85 | 4.971 | 4.7225 | 5.2195
|  | RIAP L1 | 2105EC | 291 |  |  | 2.795 | 5.394 | 5.1243 | 5.6637
|  | RIAP L2 | 2228EC | 68 |  |  | 17.321 | 7.152 | 6.7944 | 7.5096
FERRITIN, U/L | C6000-A:E2-1 | [12 empty] | 6.1978 | 5.8879 | 6.5077
|  | RIAP L3 | 2181EC | 163 |  |  | 340.861 | 4.244 | 4.0318 | 4.4562
|  | sRIAP L1 | 2179EC | 290 |  |  | 22.798 | 6.977 | 6.6281 | 7.3259
|  | PC V1 | 595763 | 44 |  |  | 144.439 | 2.045 | 1.9427 | 2.1473
|  | sRIAP L3 | 2147 EC | 325 |  |  | 311.743 | 3.752 | 3.5644 | 3.9396
|  | sRIAP L3 | 2181EC | 307 |  |  | 339.172 | 3.731 | 3.5444 | 3.9175
|  | sRIAP L1 | 2105EC | 515 |  |  | 14.33 | 8.887 | 8.4427 | 9.3314
|  | RIAP L1 | 2105EC | 133 |  |  | 15.711 | 5.935 | 5.6382 | 6.2317
|  | RIAP L3 | 2110EC | 133 |  |  | 326.623 | 3.723 | 3.5368 | 3.9091
FOLLICLE STIMULATING HORMONE, U/L | C6000-A:E1-1 | [12 empty] | 3.2759 | 3.1121 | 3.4396
|  | RIAP L1 | 2105EC | 297 |  |  | 6.689 | 3.225 | 3.0638 | 3.3862
|  | RIAP L1 | 2179EC | 68 |  |  | 5.974 | 3.474 | 3.3003 | 3.6477
|  | RIAP L3 | 2147 | 148 |  |  | 54.04 | 2.743 | 2.6059 | 2.8801
|  | RIAP L3 | 2181EC | 71 |  |  | 52.477 | 2.777 | 2.6381 | 2.9158
|  | RIAP L3 | 2110EC | 140 |  |  | 56.794 | 4.331 | 4.1145 | 4.5476
FOLLICLE STIMULATING HORMONE, U/L | C6000-A:E1-2 | [12 empty] | 3.2759 | 3.1121 | 3.4396
|  | RIAP L3 | 2181EC | 61 |  |  | 56.469 | 2.516 | 2.3902 | 2.6418
|  | RIAP L1 | 2179EC | 62 |  |  | 6.502 | 2.846 | 2.7037 | 2.9883
FREE PSA, U/L | C6000-A:E1-2 | [12 empty] | 4.891 | 4.6465 | 5.1356
|  | RIAP L1 | 2179EC | 114 |  |  | 1.154 | 3.884 | 3.6898 | 4.0782
|  | RIAP L2 | 2228EC | 60 |  |  | 9.151 | 2.554 | 2.4263 | 2.6817
|  | RIAP L2 | 2107EC | 290 |  |  | 7.716 | 5.567 | 5.2886 | 5.8454
|  | RIAP L2 | 2124EC | 75 |  |  | 10.005 | 2.861 | 2.718 | 3.004
|  | RIAP L1 | 2105EC | 302 |  |  | 1.225 | 5.266 | 5.0027 | 5.5293
Free T4 Gen 3, U/L | C6000-A:E1-1 | [12 empty] | 3.8991 | 3.7041 | 4.094
|  | sRIAP L2 | 2228EC | 126 |  |  | 29.033 | 4.191 | 3.9815 | 4.4005
|  | RIAP L2 | 2107EC | 144 |  |  | 28.582 | 4.649 | 4.4165 | 4.8815
|  | sRIAP L1 | 2105EC | 476 |  |  | 13.497 | 3.737 | 3.5501 | 3.9238
|  | sRIAP L2 | 2107EC | 434 |  |  | 29.116 | 4.008 | 3.8076 | 4.2084
|  | sRIAP L2 | 2124EC | 147 |  |  | 29.66 | 2.48 | 2.356 | 2.604
|  | sRIAP L1 | 2179EC | 250 |  |  | 13.597 | 4.209 | 3.9985 | 4.4194
|  | RIAP L1 | 2105EC | 131 |  |  | 13.084 | 3.59 | 3.4105 | 3.7695
GAMMA-GLUTAMYL TRANSFERASE, U/L | C6000-A:C1 | [12 empty] | 2.3619 | 2.2438 | 2.48
|  | PCCC 2 | 519198 | 349 |  |  | 220.934 | 1.616 | 1.5352 | 1.6968
|  | s-PCCC 1 | 525027 | 52 |  |  | 55.212 | 3.447 | 3.2746 | 3.6193
|  | s-PCCC 2 | 461597 | 45 |  |  | 271.067 | 3.772 | 3.5834 | 3.9606
|  | PCCC 1 | 525027 | 311 |  |  | 55.1 | 2.087 | 1.9827 | 2.1913
|  | PCCC 1 | 525205 | 142 |  |  | 57.408 | 2.087 | 1.9827 | 2.1913
|  | PCCC 2 | 535719 | 90 |  |  | 226.7 | 1.198 | 1.1381 | 1.2579
GAMMA-GLUTAMYL TRANSFERASE, U/L | C6000-B:C2 | [12 empty] | 2.3619 | 2.2438 | 2.48
|  | s-PCCC 2-B | 461597 | 40 |  |  | 274.125 | 0.819 | 0.77805 | 0.85995
|  | PCCC 2-B | 519198 | 337 |  |  | 221.421 | 2.999 | 2.8491 | 3.149
|  | PCCC 1-B | 525205 | 144 |  |  | 57.014 | 2.095 | 1.9903 | 2.1997
|  | s-PCCC 1-B | 525027 | 47 |  |  | 55.532 | 1.397 | 1.3272 | 1.4668
|  | PCCC 1-B | 525027 | 291 |  |  | 54.784 | 2.958 | 2.8101 | 3.1059
|  | PCCC 2-B | 535719 | 88 |  |  | 224.591 | 1.507 | 1.4317 | 1.5823
GLUCOSE, U/L | C6000-A:C1 | [12 empty] | 2.2216 | 2.1105 | 2.3327
|  | PCCC 1 | 525205 | 130 |  |  | 5.787 | 2.197 | 2.0871 | 2.3068
|  | PCCC 2 | 535719 | 82 |  |  | 13.45 | 2.27 | 2.1565 | 2.3835
|  | PCCC 2 | 519198 | 326 |  |  | 13.47 | 2.056 | 1.9532 | 2.1588
|  | PCCC 2 | 461597 | 66 |  |  | 13.1 | 1.562 | 1.4839 | 1.6401
|  | PCCC 1 | 525027 | 360 |  |  | 5.759 | 2.048 | 1.9456 | 2.1504
GLUCOSE, U/L | C6000-B:C2 | [12 empty] | 2.2216 | 2.1105 | 2.3327
|  | PCCC 1-B | 525205 | 132 |  |  | 5.769 | 2.927 | 2.7807 | 3.0734
|  | PCCC 2-B | 519198 | 311 |  |  | 13.393 | 2.562 | 2.4339 | 2.6901
|  | PCCC 2-B | 461597 | 62 |  |  | 12.974 | 1.572 | 1.4934 | 1.6506
|  | PCCC 2-B | 535719 | 81 |  |  | 13.277 | 2.14 | 2.033 | 2.247
|  | PCCC 1-B | 525027 | 337 |  |  | 5.701 | 2.121 | 2.0149 | 2.2271
GLYCATED HAEMOGLOBIN (NGSP), U/L | C6000-A:C1 | [12 empty] | 1.7988 | 1.7088 | 1.8887
|  | PCA1N-VER | 617428 | 139 |  |  | 5.737 | 2.059 | 1.9561 | 2.162
|  | PCA1P | 617429 | 91 |  |  | 9.646 | 1.709 | 1.6236 | 1.7945
|  | PCA1N-A | 563450 | 190 |  |  | 5.794 | 2.552 | 2.4244 | 2.6796
|  | PCA1P-A | 563451 | 191 |  |  | 10.587 | 2.393 | 2.2734 | 2.5126
GLYCATED HAEMOGLOBIN (NGSP), U/L | C6000-B:C2 | [12 empty] | 1.7988 | 1.7088 | 1.8887
|  | PCA1N | 563450 | 63 |  |  | 5.737 | 1.213 | 1.1523 | 1.2736
|  | PCA1P | 563451 | 589 |  |  | 10.703 | 1.508 | 1.4326 | 1.5834
|  | PCA1P | 617429 | 245 |  |  | 9.737 | 1.76 | 1.672 | 1.848
|  | PCA1N | 509111 | 110 |  |  | 5.76 | 1.29 | 1.2255 | 1.3545
|  | PCA1N | 617428 | 357 |  |  | 5.761 | 1.891 | 1.7965 | 1.9855
|  | PCA1N | 548925 | 422 |  |  | 5.731 | 1.64 | 1.558 | 1.722
|  | PCA1P | 509113 | 110 |  |  | 10.86 | 1.126 | 1.0697 | 1.1823
HAPTOGLOBIN, U/L | C6000-B:C2 | [12 empty] | 2.5739 | 2.4452 | 2.7026
|  | s-PCCC 2-B | 461597 | 36 |  |  | 1.308 | 1.381 | 1.3119 | 1.4501
|  | s-PCCC 1-B | 525027 | 43 |  |  | 0.765 | 1.216 | 1.1552 | 1.2768
|  | PCCC 2-B | 519198 | 332 |  |  | 1.221 | 2.94 | 2.793 | 3.087
|  | PCCC 1-B | 525205 | 156 |  |  | 0.8 | 2.256 | 2.1432 | 2.3688
|  | PCCC 2-B | 535719 | 104 |  |  | 1.253 | 1.312 | 1.2464 | 1.3776
|  | PCCC 1-B | 525027 | 284 |  |  | 0.782 | 2.87 | 2.7265 | 3.0135
HDL CHOLESTEROL, U/L | C6000-A:C1 | [12 empty] | 3.4043 | 3.2341 | 3.5745
|  | PCCC 1 | 525027 | 283 |  |  | 0.764 | 3.124 | 2.9678 | 3.2802
|  | PCCC 2 | 519198 | 316 |  |  | 1.457 | 2.962 | 2.8139 | 3.1101
|  | PCCC 2 | 535719 | 84 |  |  | 1.595 | 4.044 | 3.8418 | 4.2462
|  | s-PCCC 2 | 461597 | 32 |  |  | 1.431 | 1.527 | 1.4506 | 1.6034
|  | PCCC 1 | 525205 | 131 |  |  | 0.753 | 4.857 | 4.6142 | 5.0998
|  | s-PCCC 1 | 525027 | 39 |  |  | 0.784 | 2.145 | 2.0377 | 2.2523
HIGH-SENSITIVITY TROPONIN T, U/L | C6000-A:E2-1 | [12 empty] | 6.0199 | 5.7189 | 6.3209
|  | Ran Trop T | 4607CK | 198 |  |  | 6.474 | 10.625 | 10.0938 | 11.1562
|  | PC TN1 | 590216 | 163 |  |  | 27.711 | 3.068 | 2.9146 | 3.2214
|  | BIO CTC1 | 89041 | 160 |  |  | 18.187 | 3.927 | 3.7306 | 4.1234
|  | PC TN1 | 558106 | 82 |  |  | 25.588 | 2.724 | 2.5878 | 2.8602
|  | PC TN1 | 629183 | 164 |  |  | 28.741 | 3.304 | 3.1388 | 3.4692
|  | BIO CTC1 | 89021 | 53 |  |  | 18.186 | 2.653 | 2.5204 | 2.7856
HIGH-SENSITIVITY TROPONIN T, U/L | C6000-B:E3-1 | [12 empty] | 6.0199 | 5.7189 | 6.3209
|  | PC TN1-B | 590216 | 157 |  |  | 27.803 | 5.887 | 5.5926 | 6.1814
|  | PC TN1-B | 558106 | 86 |  |  | 24.246 | 2.047 | 1.9446 | 2.1494
|  | BIO CTC1-B | 89041 | 181 |  |  | 18.932 | 4.309 | 4.0935 | 4.5245
|  | Ran Trop T-B | 4607CK | 197 |  |  | 6.285 | 9.793 | 9.3033 | 10.2827
|  | PC TN1-B | 629183 | 176 |  |  | 29.951 | 4.046 | 3.8437 | 4.2483
|  | BIO CTC1-B | 89021 | 49 |  |  | 17.173 | 2.098 | 1.9931 | 2.2029
Human growth hormone, U/L | LIAISON XL | [12 empty] | 3.7499 | 3.5624 | 3.9374
|  | hGH L2 | 9260310 | 52 |  |  | 17.45 | 3.568 | 3.3896 | 3.7464
|  | hGH L1 | 7980310 | 55 |  |  | 3.771 | 3.914 | 3.7183 | 4.1097
IGG CSF, U/L | C6000-A:C1 | [12 empty] | 4.2849 | 4.0707 | 4.4992
|  | PNPUC TPC3 | 649811 | 157 |  |  | 21.281 | 4.28 | 4.066 | 4.494
|  | PNPUC TPC3 | 531179 | 75 |  |  | 20.031 | 5.101 | 4.8459 | 5.3561
|  | PPPUC TPC3 | 531181 | 32 |  |  | 134.157 | 3.374 | 3.2053 | 3.5427
|  | PNPUC TPC3 | 602712 | 109 |  |  | 23.003 | 6.289 | 5.9745 | 6.6034
|  | PNPUC TPC3 | 673292 | 35 |  |  | 19.981 | 4.753 | 4.5153 | 4.9906
|  | PPPUC TPC3 | 602710 | 154 |  |  | 133.509 | 3.278 | 3.1141 | 3.4419
|  | PPPUC TPC3 | 637091 | 101 |  |  | 142.362 | 2.534 | 2.4073 | 2.6607
|  | PPPUC TPC3 | 646927 | 112 |  |  | 128.098 | 3.821 | 3.63 | 4.0121
IMMUNOGLOBULIN A, U/L | C6000-B:C2 | [12 empty] | 2.5071 | 2.3817 | 2.6324
|  | PCCC 2-B | 535719 | 63 |  |  | 2.234 | 1.252 | 1.1894 | 1.3146
|  | PCCC 1-B | 525027 | 195 |  |  | 1.538 | 2.907 | 2.7616 | 3.0524
|  | PCCC 1-B | 525205 | 103 |  |  | 1.498 | 2.443 | 2.3209 | 2.5652
|  | PCCC 2-B | 519198 | 228 |  |  | 2.458 | 2.422 | 2.3009 | 2.5431
IMMUNOGLOBULIN G, U/L | C6000-B:C2 | [12 empty] | 2.7191 | 2.5832 | 2.8551
|  | PCCC 2-B | 519198 | 226 |  |  | 11.94 | 2.662 | 2.5289 | 2.7951
|  | PCCC 1-B | 525205 | 108 |  |  | 7.688 | 2.515 | 2.3893 | 2.6408
|  | PCCC 2-B | 535719 | 68 |  |  | 11.649 | 2.352 | 2.2344 | 2.4696
|  | PCCC 1-B | 525027 | 193 |  |  | 7.382 | 2.878 | 2.7341 | 3.0219
|  | s-PCCC 1-B | 525027 | 32 |  |  | 7.437 | 3.439 | 3.267 | 3.6109
IMMUNOGLOBULIN M, U/L | C6000-B:C2 | [12 empty] | 4.0695 | 3.866 | 4.273
|  | PCCC 1-B | 525205 | 109 |  |  | 0.673 | 4.026 | 3.8247 | 4.2273
|  | PCCC 2-B | 535719 | 67 |  |  | 1.059 | 2.29 | 2.1755 | 2.4045
|  | PCCC 2-B | 519198 | 236 |  |  | 1.09 | 3.663 | 3.4798 | 3.8462
|  | PCCC 1-B | 525027 | 199 |  |  | 0.663 | 4.925 | 4.6787 | 5.1712
INORGANIC PHOSPHATE URINE, U/L | C6000-B:C2 | [12 empty] | 2.2636 | 2.1504 | 2.3768
|  | Ran-Liq 2 | 1209UC | 213 |  |  | 8.734 | 2.433 | 2.3114 | 2.5547
|  | Bio-liq1 | 88161 | 113 |  |  | 8.302 | 2.051 | 1.9485 | 2.1536
|  | Bio-liq 2 | 88162 | 107 |  |  | 16.425 | 2.364 | 2.2458 | 2.4822
|  | Ran-Liq 3 | 1238UC | 127 |  |  | 14.63 | 2.297 | 2.1822 | 2.4118
|  | Ran-Liq 3 | 1210UC | 214 |  |  | 14.385 | 2.232 | 2.1204 | 2.3436
|  | Ran-Liq 2 | 1237UC | 115 |  |  | 8.722 | 2.052 | 1.9494 | 2.1546
INSULIN, U/L | C6000-A:E1-2 | [12 empty] | 5.7354 | 5.4486 | 6.0222
|  | RIAP L3 | 2110EC | 156 |  |  | 27.819 | 5.737 | 5.4501 | 6.0239
|  | RIAP L3 | 2147 | 164 |  |  | 26.598 | 2.509 | 2.3836 | 2.6345
|  | RIAP L3 | 2181EC | 165 |  |  | 28.213 | 4.514 | 4.2883 | 4.7397
|  | RIAP L2 | 2228EC | 81 |  |  | 15.966 | 7.071 | 6.7174 | 7.4245
|  | RIAP L2 | 2124EC | 98 |  |  | 11.855 | 10.657 | 10.1242 | 11.1898
|  | RIAP L2 | 2107EC | 316 |  |  | 13.926 | 4.897 | 4.6521 | 5.1419
IRON, U/L | C6000-A:C1 | [12 empty] | 2.5802 | 2.4512 | 2.7092
|  | PCCC 2 | 461597 | 67 |  |  | 43.791 | 1.625 | 1.5437 | 1.7063
|  | PCCC 2 | 535719 | 93 |  |  | 43.69 | 1.794 | 1.7043 | 1.8837
|  | PCCC 2 | 519198 | 375 |  |  | 43.156 | 2.031 | 1.9295 | 2.1326
|  | PCCC 1 | 525205 | 152 |  |  | 19.509 | 2.404 | 2.2838 | 2.5242
|  | PCCC 1 | 525027 | 401 |  |  | 19.892 | 3.288 | 3.1236 | 3.4524
Insulin-like growth factor I, U/L | LIAISON XL | [12 empty] | 15.8222 | 15.0311 | 16.6134
|  | IGF-1 L2 | 1422081 | 36 |  |  | 233.297 | 17.711 | 16.8255 | 18.5966
|  | IGF-1 L2 | 1422084 | 37 |  |  | 238.219 | 15.914 | 15.1183 | 16.7097
|  | IGF-1 L1 | 1412084 | 35 |  |  | 60.231 | 13.493 | 12.8184 | 14.1677
LACTATE, U/L | C6000-B:C2 | [12 empty] | 2.6957 | 2.5609 | 2.8305
|  | PCCC 1-B | 525027 | 231 |  |  | 1.76 | 2.952 | 2.8044 | 3.0996
|  | PCCC 2-B | 535719 | 83 |  |  | 3.811 | 1.474 | 1.4003 | 1.5477
|  | s-PCCC 1-B | 525027 | 33 |  |  | 1.754 | 1.548 | 1.4706 | 1.6254
|  | PCCC 1-B | 525205 | 135 |  |  | 1.781 | 3.131 | 2.9745 | 3.2875
|  | s-PCCC 2-B | 461597 | 30 |  |  | 3.658 | 0.942 | 0.8949 | 0.9891
|  | PCCC 2-B | 519198 | 271 |  |  | 3.818 | 2.754 | 2.6163 | 2.8917
LACTATE DEHYDROGENASE, U/L | C6000-A:C1 | [12 empty] | 2.6682 | 2.5347 | 2.8016
|  | s-PCCC 1 | 525027 | 47 |  |  | 175.106 | 1.137 | 1.0801 | 1.1939
|  | PCCC 2 | 519198 | 361 |  |  | 308.645 | 2.192 | 2.0824 | 2.3016
|  | PCCC 2 | 535719 | 98 |  |  | 301.398 | 2.295 | 2.1803 | 2.4097
|  | s-PCCC 2 | 461597 | 40 |  |  | 310.375 | 1.036 | 0.9842 | 1.0878
|  | PCCC 1 | 525027 | 321 |  |  | 173.555 | 3.006 | 2.8557 | 3.1563
|  | PCCC 1 | 525205 | 155 |  |  | 165.316 | 2.336 | 2.2192 | 2.4528
LACTATE DEHYDROGENASE, U/L | C6000-B:C2 | [12 empty] | 2.6682 | 2.5347 | 2.8016
|  | PCCC 2-B | 535719 | 88 |  |  | 297.455 | 2.529 | 2.4026 | 2.6555
|  | PCCC 2-B | 519198 | 295 |  |  | 312.736 | 3.101 | 2.9459 | 3.2561
|  | PCCC 1-B | 525027 | 256 |  |  | 175.336 | 2.555 | 2.4272 | 2.6827
|  | PCCC 1-B | 525205 | 138 |  |  | 166.159 | 3.786 | 3.5967 | 3.9753
|  | s-PCCC 1-B | 525027 | 39 |  |  | 174.769 | 1.528 | 1.4516 | 1.6044
|  | s-PCCC 2-B | 461597 | 34 |  |  | 308.882 | 1.423 | 1.3518 | 1.4942
LIPASE, U/L | C6000-A:C1 | [12 empty] | 3.3757 | 3.2069 | 3.5445
|  | s-PCCC 2 | 461597 | 43 |  |  | 93.747 | 3.123 | 2.9668 | 3.2792
|  | PCCC 2 | 535719 | 100 |  |  | 97.478 | 3.128 | 2.9716 | 3.2844
|  | PCCC 1 | 525027 | 338 |  |  | 44.378 | 3.547 | 3.3697 | 3.7243
|  | s-PCCC 1 | 525027 | 47 |  |  | 43.721 | 2.308 | 2.1926 | 2.4234
|  | PCCC 2 | 519198 | 382 |  |  | 100.585 | 3.558 | 3.3801 | 3.7359
|  | PCCC 1 | 525205 | 161 |  |  | 44.065 | 3.01 | 2.8595 | 3.1605
LIPOPROTEIN A2, U/L | C6000-B:C2 | [12 empty] | 5.5129 | 5.2372 | 5.7885
|  | PC LPA 1 | 520719 | 95 |  |  | 42.46 | 4.809 | 4.5686 | 5.0495
|  | PC LPA 2 | 520720 | 92 |  |  | 108.754 | 6.156 | 5.8482 | 6.4638
LITHIUM, U/L | C6000-B:C2 | [12 empty] | 3.769 | 3.5806 | 3.9575
|  | s-PCCC 1-B | 525027 | 37 |  |  | 0.939 | 2.613 | 2.4823 | 2.7437
|  | PCCC 2-B | 519198 | 276 |  |  | 1.794 | 3.096 | 2.9412 | 3.2508
|  | PCCC 1-B | 525205 | 135 |  |  | 0.919 | 4.362 | 4.1439 | 4.5801
|  | s-PCCC 2-B | 461597 | 34 |  |  | 1.796 | 2.554 | 2.4263 | 2.6817
|  | PCCC 2-B | 535719 | 85 |  |  | 1.795 | 3.705 | 3.5198 | 3.8902
|  | PCCC 1-B | 525027 | 243 |  |  | 0.908 | 4.37 | 4.1515 | 4.5885
LUTEINISING HORMONE, U/L | C6000-A:E1-2 | [12 empty] | 4.2035 | 3.9933 | 4.4137
|  | RIAP L3 | 2181EC | 132 |  |  | 54.598 | 4.678 | 4.4441 | 4.9119
|  | RIAP L3 | 2110EC | 136 |  |  | 53.477 | 3.993 | 3.7934 | 4.1927
|  | RIAP L1 | 2105EC | 298 |  |  | 3.576 | 3.565 | 3.3868 | 3.7433
|  | RIAP L1 | 2179EC | 125 |  |  | 3.748 | 5.852 | 5.5594 | 6.1446
|  | RIAP L3 | 2147 | 154 |  |  | 55.878 | 3.432 | 3.2604 | 3.6036
MAGNESIUM, U/L | C6000-A:C1 | [12 empty] | 2.5527 | 2.4251 | 2.6804
|  | PCCC 2 | 461597 | 71 |  |  | 1.283 | 1.348 | 1.2806 | 1.4154
|  | PCCC 1 | 525027 | 417 |  |  | 0.792 | 2.373 | 2.2544 | 2.4916
|  | PCCC 1 | 525205 | 165 |  |  | 0.79 | 2.317 | 2.2012 | 2.4329
|  | PCCC 2 | 535719 | 103 |  |  | 1.348 | 2.167 | 2.0587 | 2.2753
|  | PCCC 2 | 519198 | 389 |  |  | 1.35 | 2.091 | 1.9865 | 2.1955
MAGNESIUM, U/L | C6000-B:C2 | [12 empty] | 2.5527 | 2.4251 | 2.6804
|  | PCCC 2-B | 535719 | 99 |  |  | 1.346 | 3.716 | 3.5302 | 3.9018
|  | PCCC 1-B | 525205 | 168 |  |  | 0.79 | 3.907 | 3.7117 | 4.1024
|  | PCCC 2-B | 461597 | 71 |  |  | 1.294 | 1.677 | 1.5932 | 1.7609
|  | PCCC 1-B | 525027 | 388 |  |  | 0.794 | 2.725 | 2.5888 | 2.8613
|  | PCCC 2-B | 519198 | 370 |  |  | 1.355 | 2.351 | 2.2334 | 2.4686
MAGNESIUM URINE, U/L | C6000-B:C2 | [12 empty] | 7.1869 | 6.8276 | 7.5463
|  | Ran-Liq 2 | 1209UC | 215 |  |  | 2.63 | 11.858 | 11.2651 | 12.4509
|  | Ran-Liq 3 | 1210UC | 216 |  |  | 5.35 | 6.355 | 6.0373 | 6.6728
|  | Ran-Liq 3 | 1238UC | 134 |  |  | 6.415 | 3.342 | 3.1749 | 3.5091
|  | Bio-liq 2 | 88162 | 125 |  |  | 3.925 | 2.894 | 2.7493 | 3.0387
|  | Ran-Liq 2 | 1237UC | 122 |  |  | 3.091 | 3.845 | 3.6528 | 4.0373
|  | Bio-liq1 | 88161 | 126 |  |  | 1.652 | 6.358 | 6.0401 | 6.6759
OESTRADIOL III, U/L | C6000-A:E1-1 | [12 empty] | 7.5589 | 7.181 | 7.9369
|  | RIAP L1 | 2179EC | 130 |  |  | 201.844 | 7.724 | 7.3378 | 8.1102
|  | RIAP L3 | 2110EC | 143 |  |  | 1245.881 | 3.8 | 3.61 | 3.99
|  | RIAP L3 | 2147 | 157 |  |  | 1507.191 | 3.164 | 3.0058 | 3.3222
|  | RIAP L3 | 2181EC | 133 |  |  | 1512.406 | 2.364 | 2.2458 | 2.4822
|  | RIAP L1 | 2105EC | 303 |  |  | 176.634 | 11.091 | 10.5365 | 11.6456
Osmolality, U/L | OSMO1 | [12 empty] | 0.90118 | 0.85612 | 0.94624
|  | Serim BIOLIQ 1 | 45871 | 240 |  |  | 305.983 | 0.908 | 0.8626 | 0.9534
|  | Serim BIOLIQ 1 | 45931 | 114 |  |  | 311.789 | 0.898 | 0.8531 | 0.9429
|  | Serim BIOLIQ 2 | 45952 | 74 |  |  | 400.23 | 1.146 | 1.0887 | 1.2033
|  | Serim BIOLIQ 2 | 45872 | 110 |  |  | 408.373 | 0.783 | 0.74385 | 0.82215
|  | Serim BIOLIQ 2 | 45912 | 139 |  |  | 411.475 | 0.863 | 0.81985 | 0.90615
|  | Serim BIOLIQ 2 | 45862 | 41 |  |  | 401.244 | 0.784 | 0.7448 | 0.8232
PARACETAMOL 2, U/L | C6000-B:C2 | [12 empty] | 4.8694 | 4.626 | 5.1129
|  | RIAP L2-B | 2107EC | 280 |  |  | 319.991 | 4.752 | 4.5144 | 4.9896
|  | RIAP L2-B | 2228EC | 70 |  |  | 324.084 | 6.388 | 6.0686 | 6.7074
|  | RIAP L1-B | 2105EC | 288 |  |  | 152.374 | 4.181 | 3.972 | 4.39
|  | RIAP L2-B | 2124EC | 77 |  |  | 322.526 | 4.565 | 4.3368 | 4.7932
|  | RIAP L1-B | 2179EC | 123 |  |  | 154.079 | 5.751 | 5.4634 | 6.0386
PARATHYROID HORMONE, U/L | C6000-A:E2-1 | [12 empty] | 3.651 | 3.4684 | 3.8335
|  | PC V2 | 581988 | 55 |  |  | 20.859 | 1.373 | 1.3043 | 1.4417
|  | PC V1 | 612148 | 66 |  |  | 6.367 | 3.088 | 2.9336 | 3.2424
|  | PC V1 | 581987 | 55 |  |  | 6.467 | 4.138 | 3.9311 | 4.3449
|  | PC V2 | 573356 | 92 |  |  | 20.016 | 4.076 | 3.8722 | 4.2798
|  | PC V1 | 573355 | 92 |  |  | 6.138 | 4.493 | 4.2683 | 4.7177
|  | PC V1 | 595763 | 209 |  |  | 6.668 | 2.937 | 2.7902 | 3.0838
|  | PC V2 | 595765 | 235 |  |  | 19.985 | 4.048 | 3.8456 | 4.2504
PHOSPHATE, U/L | C6000-A:C1 | [12 empty] | 2.2882 | 2.1737 | 2.4026
|  | PCCC 2 | 461597 | 72 |  |  | 2.521 | 1.531 | 1.4545 | 1.6076
|  | PCCC 1 | 525027 | 417 |  |  | 1.29 | 2.974 | 2.8253 | 3.1227
|  | PCCC 2 | 519198 | 389 |  |  | 2.544 | 1.909 | 1.8135 | 2.0044
|  | PCCC 1 | 525205 | 164 |  |  | 1.159 | 2.05 | 1.9475 | 2.1525
|  | PCCC 2 | 535719 | 104 |  |  | 2.691 | 1.271 | 1.2074 | 1.3345
PHOSPHATE, U/L | C6000-B:C2 | [12 empty] | 2.2882 | 2.1737 | 2.4026
|  | PCCC 2-B | 535719 | 89 |  |  | 2.721 | 1.968 | 1.8696 | 2.0664
|  | PCCC 1-B | 525027 | 366 |  |  | 1.288 | 1.988 | 1.8886 | 2.0874
|  | PCCC 2-B | 519198 | 346 |  |  | 2.564 | 2.281 | 2.1669 | 2.395
|  | PCCC 1-B | 525205 | 151 |  |  | 1.19 | 3.172 | 3.0134 | 3.3306
|  | PCCC 2-B | 461597 | 70 |  |  | 2.464 | 1.202 | 1.1419 | 1.2621
POTASSIUM, U/L | C6000-A:ISE1 | [12 empty] | 1.7767 | 1.6879 | 1.8655
|  | s-PCCC 2 | 535719 | 192 |  |  | 7.393 | 1.631 | 1.5494 | 1.7126
|  | PCCC 1 | 525027 | 106 |  |  | 3.65 | 1.654 | 1.5713 | 1.7367
|  | PCCC 2 | 461597 | 88 |  |  | 6.833 | 1.19 | 1.1305 | 1.2495
|  | s-PCCC 1 | 525205 | 330 |  |  | 3.586 | 2.236 | 2.1242 | 2.3478
|  | s-PCCC 2 | 519198 | 717 |  |  | 7.149 | 1.669 | 1.5856 | 1.7525
|  | s-PCCC 1 | 525027 | 597 |  |  | 3.645 | 1.915 | 1.8193 | 2.0107
POTASSIUM, U/L | C6000-B:ISE1 | [12 empty] | 1.7767 | 1.6879 | 1.8655
|  | s-PCCC 2-B | 535719 | 190 |  |  | 7.306 | 1.328 | 1.2616 | 1.3944
|  | PCCC 2-B | 461597 | 92 |  |  | 6.841 | 0.81 | 0.7695 | 0.8505
|  | s-PCCC 1-B | 525027 | 565 |  |  | 3.639 | 1.833 | 1.7413 | 1.9246
|  | PCCC 1-B | 525027 | 368 |  |  | 3.654 | 1.601 | 1.521 | 1.681
|  | PCCC 2-B | 519198 | 280 |  |  | 7.125 | 1.289 | 1.2246 | 1.3535
|  | s-PCCC 2-B | 519198 | 686 |  |  | 7.113 | 1.56 | 1.482 | 1.638
|  | s-PCCC 1-B | 525205 | 293 |  |  | 3.547 | 2.665 | 2.5318 | 2.7982
POTASSIUM URINE, U/L | C6000-B:ISE1 | [12 empty] | 3.7207 | 3.5347 | 3.9068
|  | Ran-Liq 2 | 1209UC | 247 |  |  | 29.103 | 3.281 | 3.117 | 3.4451
|  | Ran-Liq 3 | 1210UC | 249 |  |  | 71.532 | 5.462 | 5.1889 | 5.7351
|  | Bio-liq 2 | 88162 | 137 |  |  | 67.437 | 2.765 | 2.6267 | 2.9032
|  | Bio-liq1 | 88161 | 144 |  |  | 30.823 | 1.968 | 1.8696 | 2.0664
|  | Ran-Liq 2 | 1237UC | 137 |  |  | 29.08 | 2.725 | 2.5888 | 2.8613
|  | Ran-Liq 3 | 1238UC | 149 |  |  | 72.903 | 3.66 | 3.477 | 3.843
PROCALCITONIN (SENSITIVE), U/L | C6000-A:E1-2 | [12 empty] | 3.6456 | 3.4633 | 3.8278
|  | PC PCT1 | 577640 | 33 |  |  | 0.462 | 3.161 | 3.0029 | 3.319
|  | PC PCT2 | 577646 | 33 |  |  | 8.808 | 4.033 | 3.8314 | 4.2347
PROCALCITONIN (SENSITIVE), U/L | C6000-A:E2-1 | [12 empty] | 3.6456 | 3.4633 | 3.8278
|  | PC PCT2 | 671599 | 107 |  |  | 9.08 | 2.113 | 2.0074 | 2.2186
|  | PC PCT2 | 614523 | 61 |  |  | 9.569 | 2.66 | 2.527 | 2.793
|  | PC PCT1 | 671597 | 108 |  |  | 0.491 | 2.904 | 2.7588 | 3.0492
|  | PC PCT1 | 614519 | 61 |  |  | 0.496 | 2.425 | 2.3037 | 2.5463
|  | PC PCT2 | 559763 | 182 |  |  | 9.419 | 4.027 | 3.8256 | 4.2284
|  | PC PCT1 | 559760 | 181 |  |  | 0.509 | 4.799 | 4.5591 | 5.039
PROGESTERONE III, U/L | C6000-A:E2-2 | [12 empty] | 6.5456 | 6.2184 | 6.8729
|  | RIAP L1 | 2105EC | 287 |  |  | 3.378 | 6.885 | 6.5407 | 7.2293
|  | RIAP L3 | 2181EC | 135 |  |  | 98.292 | 4.296 | 4.0812 | 4.5108
|  | RIAP L3 | 2147 | 148 |  |  | 95.094 | 3.648 | 3.4656 | 3.8304
|  | RIAP L3 | 2110EC | 136 |  |  | 89.168 | 2.824 | 2.6828 | 2.9652
|  | RIAP L1 | 2179EC | 123 |  |  | 3.357 | 11.538 | 10.9611 | 12.1149
PROLACTIN, U/L | C6000-A:E1-2 | [12 empty] | 5.0234 | 4.7722 | 5.2746
|  | RIAP L3 | 2110EC | 137 |  |  | 51.014 | 3.153 | 2.9954 | 3.3106
|  | RIAP L3 | 2181EC | 133 |  |  | 69.197 | 5.708 | 5.4226 | 5.9934
|  | RIAP L3 | 2147 | 153 |  |  | 44.917 | 5.06 | 4.807 | 5.313
|  | RIAP L2 | 2107EC | 279 |  |  | 28.545 | 4.706 | 4.4707 | 4.9413
|  | RIAP L2 | 2124EC | 78 |  |  | 26.833 | 8.272 | 7.8584 | 8.6856
|  | RIAP L2 | 2228EC | 64 |  |  | 32.167 | 1.674 | 1.5903 | 1.7577
PROSTATE-SPECIFIC AG, U/L | C6000-A:E2-2 | [12 empty] | 6.3552 | 6.0374 | 6.6729
|  | RIAP L2 | 2124EC | 99 |  |  | 12.504 | 3.75 | 3.5625 | 3.9375
|  | RIAP L2 | 2107EC | 340 |  |  | 9.828 | 7.41 | 7.0395 | 7.7805
|  | RIAP L2 | 2228EC | 77 |  |  | 12.125 | 2.93 | 2.7835 | 3.0765
|  | RIAP L1 | 2105EC | 350 |  |  | 1.584 | 7.166 | 6.8077 | 7.5243
|  | RIAP L1 | 2179EC | 146 |  |  | 1.588 | 3.81 | 3.6195 | 4.0005
PROTEIN CSF, U/L | C6000-A:C1 | [12 empty] | 3.43 | 3.2585 | 3.6015
|  | PNPUC TPC3 | 673292 | 41 |  |  | 0.214 | 3.396 | 3.2262 | 3.5658
|  | PNPUC TPC3 | 602712 | 62 |  |  | 0.214 | 3.389 | 3.2195 | 3.5585
|  | PPPUC TPC3 | 602710 | 165 |  |  | 1.488 | 1.389 | 1.3196 | 1.4585
|  | PPPUC TPC3 | 531181 | 35 |  |  | 1.43 | 1.081 | 1.027 | 1.135
|  | PNPUC TPC3 | 531179 | 79 |  |  | 0.211 | 2.808 | 2.6676 | 2.9484
|  | PPPUC TPC3 | 646927 | 132 |  |  | 1.52 | 1.718 | 1.6321 | 1.8039
|  | PNPUC TPC3 | 634291 | 30 |  |  | 0.212 | 11.011 | 10.4604 | 11.5616
|  | PNPUC TPC3 | 649811 | 186 |  |  | 0.21 | 3.89 | 3.6955 | 4.0845
|  | PPPUC TPC3 | 637091 | 122 |  |  | 1.616 | 2.991 | 2.8415 | 3.1406
PROTEIN URINE, U/L | C6000-A:C1 | [12 empty] | 3.9337 | 3.7371 | 4.1304
|  | PP PUC | 64692700 | 92 |  |  | 1.264 | 1.92 | 1.824 | 2.016
|  | PP PUC | 637091 | 122 |  |  | 1.342 | 2.016 | 1.9152 | 2.1168
|  | PP PUC | 602710 | 162 |  |  | 1.239 | 2.046 | 1.9437 | 2.1483
|  | PN PUC | 602712 | 120 |  |  | 0.178 | 3.139 | 2.9821 | 3.2959
|  | Ran-Liq 3 | 1238UC | 51 |  |  | 0.739 | 3.91 | 3.7145 | 4.1055
|  | Ran-Liq 2 | 1237UC | 50 |  |  | 0.135 | 4.105 | 3.8998 | 4.3103
|  | Ran-Liq 2 | 1209UC | 66 |  |  | 0.137 | 7.484 | 7.1098 | 7.8582
|  | PN PUC | 531179 | 81 |  |  | 0.175 | 3.148 | 2.9906 | 3.3054
|  | Ran-Liq 3 | 1210UC | 66 |  |  | 0.769 | 3.056 | 2.9032 | 3.2088
|  | PN PUC | 649811 | 203 |  |  | 0.173 | 5.783 | 5.4939 | 6.0722
|  | PP PUC | 531181 | 38 |  |  | 1.252 | 0.951 | 0.90345 | 0.99855
RHD3, U/L | C6000-B:C2 | [12 empty] | 2.6478 | 2.5154 | 2.7802
|  | PPA1c H | 563451 | 515 |  |  | 10.612 | 3.079 | 2.9251 | 3.233
|  | PPA1c H | 509113 | 51 |  |  | 10.726 | 2.216 | 2.1052 | 2.3268
|  | PNA1c H | 509111 | 51 |  |  | 5.725 | 1.995 | 1.8953 | 2.0947
|  | PPA1c H | 617429 | 222 |  |  | 9.444 | 2.733 | 2.5964 | 2.8697
|  | PNA1c H | 617428 | 224 |  |  | 5.684 | 2.241 | 2.129 | 2.3531
|  | PNA1c H | 548925 | 344 |  |  | 5.687 | 2.539 | 2.412 | 2.666
|  | PNA1c H | 563450 | 165 |  |  | 5.755 | 2.025 | 1.9238 | 2.1263
RHEUMATOID FACTOR, U/L | C6000-B:C2 | [12 empty] | 3.9009 | 3.7059 | 4.096
|  | RFCO2 | 590051 | 116 |  |  | 52.38 | 3.15 | 2.9925 | 3.3075
|  | RFCO1 | 661133 | 131 |  |  | 23.404 | 4.107 | 3.9017 | 4.3124
|  | RFCO2 | 661134 | 134 |  |  | 53.296 | 3.727 | 3.5406 | 3.9133
|  | RFCO1 | 590050 | 119 |  |  | 22.129 | 4.996 | 4.7462 | 5.2458
|  | RFCO1 | 542412 | 33 |  |  | 21.355 | 2.306 | 2.1907 | 2.4213
|  | RFCO2 | 542413 | 33 |  |  | 50.661 | 2.651 | 2.5185 | 2.7835
Renin, U/L | LIAISON XL | [12 empty] | 5.6133 | 5.3327 | 5.894
|  | RENIN L2 | 7530320 | 42 |  |  | 111.185 | 5.659 | 5.376 | 5.942
|  | RENIN L1 | 7520320 | 49 |  |  | 24.972 | 5.574 | 5.2953 | 5.8527
SERUM FOLATE III, U/L | C6000-A:E2-2 | [12 empty] | 13.0086 | 12.3582 | 13.659
|  | RIAP L2 | 2228EC | 88 |  |  | 12.459 | 31.268 | 29.7046 | 32.8314
|  | RIAP L3 | 2147 | 198 |  |  | 23.148 | 6.955 | 6.6072 | 7.3028
|  | RIAP L3 | 2110EC | 159 |  |  | 23.348 | 5.985 | 5.6857 | 6.2843
|  | RIAP L3 | 2181EC | 174 |  |  | 22.875 | 14.869 | 14.1256 | 15.6125
|  | RIAP L2 | 2124EC | 98 |  |  | 12.508 | 14.032 | 13.3304 | 14.7336
|  | RIAP L2 | 2107EC | 352 |  |  | 12.134 | 7.969 | 7.5705 | 8.3674
SEX HORMONE BINDING GLOBULIN, U/L | C6000-A:E2-2 | [12 empty] | 4.2315 | 4.02 | 4.4431
|  | RIAP L1 | 2179EC | 115 |  |  | 9.938 | 3.383 | 3.2138 | 3.5522
|  | RIAP L1 | 2105EC | 297 |  |  | 11.022 | 5 | 4.75 | 5.25
|  | RIAP L2 | 2124EC | 82 |  |  | 38.312 | 4.717 | 4.4812 | 4.9528
|  | RIAP L2 | 2107EC | 268 |  |  | 37.77 | 3.788 | 3.5986 | 3.9774
|  | RIAP L2 | 2228EC | 66 |  |  | 35.521 | 2.476 | 2.3522 | 2.5998
SODIUM, U/L | C6000-A:ISE1 | [12 empty] | 1.7338 | 1.6471 | 1.8205
|  | s-PCCC 2 | 535719 | 193 |  |  | 138.373 | 1.367 | 1.2987 | 1.4353
|  | s-PCCC 1 | 525027 | 588 |  |  | 114.427 | 1.614 | 1.5333 | 1.6947
|  | PCCC 2 | 461597 | 88 |  |  | 133.534 | 1.362 | 1.2939 | 1.4301
|  | s-PCCC 1 | 525205 | 342 |  |  | 114.234 | 1.755 | 1.6672 | 1.8428
|  | PCCC 1 | 525027 | 106 |  |  | 113.896 | 1.57 | 1.4915 | 1.6485
|  | s-PCCC 2 | 519198 | 714 |  |  | 137.38 | 1.423 | 1.3518 | 1.4942
SODIUM, U/L | C6000-B:ISE1 | [12 empty] | 1.7338 | 1.6471 | 1.8205
|  | s-PCCC 1-B | 525205 | 302 |  |  | 113.228 | 1.899 | 1.804 | 1.9939
|  | s-PCCC 2-B | 535719 | 188 |  |  | 137.287 | 4.199 | 3.9891 | 4.4089
|  | PCCC 2-B | 519198 | 279 |  |  | 136.867 | 1.356 | 1.2882 | 1.4238
|  | PCCC 1-B | 525027 | 369 |  |  | 114.388 | 1.521 | 1.4449 | 1.5971
|  | PCCC 2-B | 461597 | 94 |  |  | 133.968 | 1.113 | 1.0574 | 1.1686
|  | s-PCCC 2-B | 519198 | 697 |  |  | 136.933 | 1.476 | 1.4022 | 1.5498
|  | s-PCCC 1-B | 525027 | 560 |  |  | 114.418 | 1.565 | 1.4868 | 1.6433
SODIUM URINE, U/L | C6000-B:ISE1 | [12 empty] | 7.6652 | 7.282 | 8.0485
|  | Ran-Liq 2 | 1209UC | 241 |  |  | 70.676 | 14.005 | 13.3048 | 14.7052
|  | Ran-Liq 3 | 1210UC | 252 |  |  | 165.329 | 7.702 | 7.3169 | 8.0871
|  | Bio-liq 2 | 88162 | 164 |  |  | 167.335 | 2.107 | 2.0017 | 2.2123
|  | Ran-Liq 3 | 1238UC | 155 |  |  | 172.542 | 1.691 | 1.6064 | 1.7755
|  | Ran-Liq 2 | 1237UC | 138 |  |  | 84.92 | 2.681 | 2.5469 | 2.815
|  | Bio-liq1 | 88161 | 165 |  |  | 77.57 | 2.528 | 2.4016 | 2.6544
TESTOSTERONE, U/L | C6000-A:E2-1 | [12 empty] | 5.965 | 5.6668 | 6.2633
|  | RIAP L3 | 2181EC | 129 |  |  | 27.49 | 3.726 | 3.5397 | 3.9123
|  | RIAP L2 | 2228EC | 63 |  |  | 14.948 | 3.181 | 3.0219 | 3.3401
|  | RIAP L3 | 2147 | 145 |  |  | 27.647 | 5.899 | 5.604 | 6.194
|  | RIAP L2 | 2107EC | 273 |  |  | 12.882 | 7.282 | 6.9179 | 7.6461
|  | RIAP L2 | 2124EC | 78 |  |  | 14.379 | 3.188 | 3.0286 | 3.3474
|  | RIAP L3 | 2110EC | 134 |  |  | 28.937 | 6.944 | 6.5968 | 7.2912
THYROGLOBULIN v2, U/L | C6000-A:E1-1 | [12 empty] | 7.1672 | 6.8088 | 7.5255
|  | RIAP L3 | 2147 | 153 |  |  | 65.571 | 4.306 | 4.0907 | 4.5213
|  | RIAP L1 | 2105EC | 302 |  |  | 2.099 | 5.159 | 4.901 | 5.4169
|  | RIAP L3 | 2181EC | 130 |  |  | 62.073 | 10.34 | 9.823 | 10.857
|  | RIAP L3 | 2110EC | 144 |  |  | 67.02 | 4.727 | 4.4906 | 4.9634
|  | RIAP L1 | 2179EC | 134 |  |  | 1.939 | 11.056 | 10.5032 | 11.6088
THYROID STIMULATING HORMONE, U/L | C6000-A:E2-2 | [12 empty] | 5.5211 | 5.245 | 5.7972
|  | RIAP L2 | 2228EC | 98 |  |  | 2.595 | 2.576 | 2.4472 | 2.7048
|  | RIAP L2 | 2107EC | 420 |  |  | 2.399 | 6.018 | 5.7171 | 6.3189
|  | RIAP L2 | 2124EC | 119 |  |  | 2.93 | 4.83 | 4.5885 | 5.0715
|  | RIAP L1 | 2179EC | 174 |  |  | 0.133 | 6.438 | 6.1161 | 6.7599
|  | RIAP L1 | 2105EC | 436 |  |  | 0.103 | 4.785 | 4.5457 | 5.0243
THYROID STIMULATING HORMONE, U/L | C6000-B:E3-1 | [12 empty] | 5.5211 | 5.245 | 5.7972
|  | RIAP L2-B | 2124EC | 90 |  |  | 2.965 | 7.416 | 7.0452 | 7.7868
|  | RIAP L1-B | 2179EC | 141 |  |  | 0.136 | 6.233 | 5.9213 | 6.5446
|  | RIAP L1-B | 2105EC | 316 |  |  | 0.106 | 5.018 | 4.7671 | 5.2689
|  | RIAP L2-B | 2107EC | 312 |  |  | 2.392 | 5.914 | 5.6183 | 6.2097
|  | RIAP L2-B | 2228EC | 77 |  |  | 2.591 | 4.283 | 4.0689 | 4.4972
TOTAL BILIRUBIN, U/L | C6000-A:C1 | [12 empty] | 4.2068 | 3.9964 | 4.4171
|  | MAS BILIRUBIN 3 | BC23063A | 73 |  |  | 321.348 | 3.499 | 3.3241 | 3.674
|  | PCCC 2 | 519198 | 417 |  |  | 60.758 | 3.822 | 3.6309 | 4.0131
|  | PCCC 2 | 535719 | 108 |  |  | 63.277 | 3.941 | 3.7439 | 4.138
|  | PCCC 2 | 461597 | 77 |  |  | 61.884 | 3.082 | 2.9279 | 3.2361
|  | PCCC 1 | 525205 | 168 |  |  | 16.461 | 4.504 | 4.2788 | 4.7292
|  | PCCC 1 | 525027 | 451 |  |  | 15.771 | 5.593 | 5.3133 | 5.8727
|  | MAS BILIRUBIN 3 | BC24033A | 414 |  |  | 346.205 | 2.757 | 2.6191 | 2.8948
TOTAL BILIRUBIN, U/L | C6000-B:C2 | [12 empty] | 4.2068 | 3.9964 | 4.4171
|  | PCCC 2-B | 519198 | 397 |  |  | 60.03 | 4.845 | 4.6027 | 5.0873
|  | PCCC 2-B | 535719 | 99 |  |  | 61.335 | 2.977 | 2.8281 | 3.1258
|  | PCCC 1-B | 525205 | 166 |  |  | 16.427 | 4.649 | 4.4165 | 4.8815
|  | PCCC 2-B | 461597 | 78 |  |  | 62.617 | 1.79 | 1.7005 | 1.8795
|  | MAS BILIRUBIN-B | BC23063A | 72 |  |  | 324.158 | 2.844 | 2.7018 | 2.9862
|  | MAS BILIRUBIN-B | BC24033A | 385 |  |  | 342.326 | 3.409 | 3.2386 | 3.5795
|  | PCCC 1-B | 525027 | 418 |  |  | 15.66 | 4.798 | 4.5581 | 5.0379
TOTAL CHOLESTEROL, U/L | C6000-A:C1 | [12 empty] | 2.1084 | 2.0029 | 2.2138
|  | PCCC 1 | 525027 | 93 |  |  | 2.378 | 2.353 | 2.2353 | 2.4707
|  | s-PCCC 2 | 519198 | 610 |  |  | 4.467 | 1.944 | 1.8468 | 2.0412
|  | s-PCCC 1 | 525205 | 264 |  |  | 2.427 | 2.244 | 2.1318 | 2.3562
|  | s-PCCC 2 | 535719 | 165 |  |  | 4.347 | 2.356 | 2.2382 | 2.4738
|  | s-PCCC 1 | 525027 | 521 |  |  | 2.424 | 2.454 | 2.3313 | 2.5767
|  | PCCC 2 | 461597 | 76 |  |  | 4.485 | 2.195 | 2.0852 | 2.3047
TOTAL CHOLESTEROL, U/L | C6000-B:C2 | [12 empty] | 2.1084 | 2.0029 | 2.2138
|  | s-PCCC 1-B | 525205 | 239 |  |  | 2.423 | 2.401 | 2.2809 | 2.521
|  | s-PCCC 2-B | 519198 | 604 |  |  | 4.49 | 1.966 | 1.8677 | 2.0643
|  | PCCC 2-B | 461597 | 64 |  |  | 4.538 | 1.289 | 1.2246 | 1.3535
|  | PCCC 1-B | 525027 | 76 |  |  | 2.399 | 1.686 | 1.6017 | 1.7703
|  | s-PCCC 2-B | 535719 | 158 |  |  | 4.31 | 1.663 | 1.5798 | 1.7462
|  | s-PCCC 1-B | 525027 | 505 |  |  | 2.432 | 1.97 | 1.8715 | 2.0685
TOTAL PROTEIN, U/L | C6000-A:C1 | [12 empty] | 1.997 | 1.8972 | 2.0969
|  | PCCC 1 | 525027 | 408 |  |  | 48.055 | 1.616 | 1.5352 | 1.6968
|  | PCCC 2 | 461597 | 73 |  |  | 80.881 | 1.34 | 1.273 | 1.407
|  | PCCC 2 | 519198 | 387 |  |  | 78.18 | 1.584 | 1.5048 | 1.6632
|  | PCCC 2 | 535719 | 99 |  |  | 75.53 | 3.091 | 2.9364 | 3.2456
|  | PCCC 1 | 525205 | 164 |  |  | 48.256 | 2.842 | 2.6999 | 2.9841
TOTAL PROTEIN, U/L | C6000-B:C2 | [12 empty] | 1.997 | 1.8972 | 2.0969
|  | s-PCCC 1-B | 525027 | 42 |  |  | 47.695 | 1.12 | 1.064 | 1.176
|  | PCCC 1-B | 525205 | 151 |  |  | 47.27 | 2.31 | 2.1945 | 2.4255
|  | PCCC 2-B | 519198 | 327 |  |  | 77.293 | 2.019 | 1.9181 | 2.1199
|  | s-PCCC 2-B | 461597 | 38 |  |  | 80.203 | 1.087 | 1.0327 | 1.1414
|  | PCCC 2-B | 535719 | 89 |  |  | 74.152 | 2.365 | 2.2468 | 2.4832
|  | PCCC 1-B | 525027 | 277 |  |  | 47.584 | 1.924 | 1.8278 | 2.0202
TRANSFERRIN, U/L | C6000-A:C1 | [12 empty] | 2.9476 | 2.8002 | 3.0949
|  | s-PCCC 1 | 525027 | 54 |  |  | 1.954 | 3.332 | 3.1654 | 3.4986
|  | s-PCCC 2 | 461597 | 48 |  |  | 3.392 | 3.989 | 3.7896 | 4.1884
|  | PCCC 2 | 535719 | 110 |  |  | 3.356 | 1.794 | 1.7043 | 1.8837
|  | PCCC 2 | 519198 | 425 |  |  | 3.446 | 3.288 | 3.1236 | 3.4524
|  | PCCC 1 | 525205 | 188 |  |  | 1.919 | 1.867 | 1.7736 | 1.9604
|  | PCCC 1 | 525027 | 365 |  |  | 2.014 | 3.04 | 2.888 | 3.192
TRI-IODO THYRONINE, U/L | C6000-A:E1-2 | [12 empty] | 4.4766 | 4.2528 | 4.7004
|  | RIAP L2 | 2228EC | 65 |  |  | 15.641 | 2.536 | 2.4092 | 2.6628
|  | RIAP L2 | 2107EC | 276 |  |  | 13.81 | 3.468 | 3.2946 | 3.6414
|  | RIAP L2 | 2124EC | 74 |  |  | 15.164 | 3.567 | 3.3887 | 3.7454
|  | RIAP L1 | 2179EC | 117 |  |  | 5.052 | 4.878 | 4.6341 | 5.1219
|  | RIAP L1 | 2105EC | 294 |  |  | 4.403 | 5.556 | 5.2782 | 5.8338
TRIGLYCERIDE, U/L | C6000-A:C1 | [12 empty] | 2.5894 | 2.4599 | 2.7189
|  | PCCC 1 | 525027 | 341 |  |  | 1.35 | 3.507 | 3.3316 | 3.6824
|  | PCCC 2 | 519198 | 385 |  |  | 2.453 | 1.935 | 1.8382 | 2.0318
|  | PCCC 2 | 535719 | 107 |  |  | 2.444 | 2.139 | 2.032 | 2.246
|  | s-PCCC 1 | 525027 | 52 |  |  | 1.341 | 2.249 | 2.1366 | 2.3615
|  | s-PCCC 2 | 461597 | 42 |  |  | 2.443 | 1.869 | 1.7755 | 1.9625
|  | PCCC 1 | 525205 | 166 |  |  | 1.343 | 2.178 | 2.0691 | 2.2869
UREA, U/L | C6000-A:C1 | [12 empty] | 2.6416 | 2.5095 | 2.7736
|  | s-PCCC 1 | 525027 | 546 |  |  | 6.708 | 2.652 | 2.5194 | 2.7846
|  | PCCC 2 | 461597 | 76 |  |  | 19.705 | 2.134 | 2.0273 | 2.2407
|  | s-PCCC 2 | 519198 | 644 |  |  | 20.406 | 2.645 | 2.5128 | 2.7772
|  | s-PCCC 2 | 535719 | 167 |  |  | 19.949 | 3.139 | 2.9821 | 3.2959
|  | s-PCCC 1 | 525205 | 267 |  |  | 6.6 | 3.154 | 2.9963 | 3.3117
|  | PCCC 1 | 525027 | 95 |  |  | 6.636 | 1.915 | 1.8193 | 2.0107
UREA, U/L | C6000-B:C2 | [12 empty] | 2.6416 | 2.5095 | 2.7736
|  | PCCC 2-B | 461597 | 77 |  |  | 19.677 | 1.878 | 1.7841 | 1.9719
|  | PCCC 1-B | 525027 | 91 |  |  | 6.56 | 1.953 | 1.8554 | 2.0507
|  | s-PCCC 1-B | 525027 | 530 |  |  | 6.778 | 2.666 | 2.5327 | 2.7993
|  | s-PCCC 1-B | 525205 | 256 |  |  | 6.739 | 2.981 | 2.8319 | 3.1301
|  | s-PCCC 2-B | 519198 | 639 |  |  | 20.671 | 2.502 | 2.3769 | 2.6271
|  | s-PCCC 2-B | 535719 | 164 |  |  | 20.051 | 2.175 | 2.0663 | 2.2837
UREA URINE, U/L | C6000-A:C1 | [12 empty] | 5.8892 | 5.5947 | 6.1837
|  | Ran-Liq 2 | 1237UC | 50 |  |  | 149.36 | 3.865 | 3.6717 | 4.0583
|  | Ran-Liq 3 | 1238UC | 49 |  |  | 281.529 | 4.269 | 4.0556 | 4.4825
|  | Ran-Liq 3 | 1210UC | 73 |  |  | 275.344 | 7.779 | 7.39 | 8.1679
|  | Ran-Liq 2 | 1209UC | 73 |  |  | 153.532 | 16.514 | 15.6883 | 17.3397
UREA URINE, U/L | C6000-B:C2 | [12 empty] | 5.8892 | 5.5947 | 6.1837
|  | Ran-Liq 2 | 1209UC | 194 |  |  | 149.709 | 3.381 | 3.2119 | 3.5501
|  | Bio-liq1 | 88161 | 136 |  |  | 151.508 | 3.45 | 3.2775 | 3.6225
|  | Ran-Liq 3 | 1210UC | 229 |  |  | 278.159 | 5.411 | 5.1405 | 5.6815
|  | Bio-liq 2 | 88162 | 131 |  |  | 257.644 | 3.843 | 3.6509 | 4.0351
|  | Ran-Liq 2 | 1237UC | 122 |  |  | 149.59 | 3.724 | 3.5378 | 3.9102
|  | Ran-Liq 3 | 1238UC | 137 |  |  | 285.408 | 3.184 | 3.0248 | 3.3432
URIC ACID, U/L | C6000-A:C1 | [12 empty] | 2.046 | 1.9437 | 2.1483
|  | PCCC 1 | 525205 | 141 |  |  | 0.287 | 2.129 | 2.0225 | 2.2355
|  | PCCC 1 | 525027 | 357 |  |  | 0.289 | 2.198 | 2.0881 | 2.3079
|  | PCCC 2 | 461597 | 67 |  |  | 0.571 | 1.574 | 1.4953 | 1.6527
|  | PCCC 2 | 519198 | 326 |  |  | 0.588 | 2.119 | 2.013 | 2.225
|  | PCCC 2 | 535719 | 93 |  |  | 0.605 | 2.011 | 1.9104 | 2.1115
URIC ACID, U/L | C6000-B:C2 | [12 empty] | 2.046 | 1.9437 | 2.1483
|  | PCCC 2-B | 535719 | 79 |  |  | 0.596 | 2.095 | 1.9903 | 2.1997
|  | PCCC 1-B | 525027 | 255 |  |  | 0.287 | 2.063 | 1.9599 | 2.1662
|  | s-PCCC 2-B | 461597 | 33 |  |  | 0.569 | 0.922 | 0.8759 | 0.9681
|  | PCCC 1-B | 525205 | 134 |  |  | 0.285 | 2.003 | 1.9028 | 2.1031
|  | PCCC 2-B | 519198 | 300 |  |  | 0.584 | 2.009 | 1.9085 | 2.1094
|  | s-PCCC 1-B | 525027 | 38 |  |  | 0.286 | 1.104 | 1.0488 | 1.1592
URIC ACID URINE, U/L | C6000-B:C2 | [12 empty] | 4.5006 | 4.2756 | 4.7257
|  | Bio-liq1 | 88161 | 104 |  |  | 0.854 | 2.474 | 2.3503 | 2.5977
|  | Bio-liq 2 | 88162 | 99 |  |  | 1.282 | 2.279 | 2.165 | 2.3929
|  | Ran-Liq 3 | 1238UC | 117 |  |  | 0.871 | 2.163 | 2.0549 | 2.2712
|  | Ran-Liq 2 | 1237UC | 108 |  |  | 0.545 | 2.637 | 2.5051 | 2.7689
|  | Ran-Liq 3 | 1210UC | 198 |  |  | 0.861 | 4.613 | 4.3824 | 4.8437
|  | Ran-Liq 2 | 1209UC | 197 |  |  | 0.541 | 7.118 | 6.7621 | 7.4739
Urine osmolality, U/L | OSMO1 | [12 empty] | 0.7054 | 0.67013 | 0.74067
|  | Biorad Urine QC1 | 88162 | 93 |  |  | 787.753 | 0.411 | 0.39045 | 0.43155
|  | Ran-Liq 2 | 1237UC | 79 |  |  | 436.139 | 1.175 | 1.1162 | 1.2337
|  | Biorad Urine QC1 | 88161 | 102 |  |  | 436.5 | 0.438 | 0.4161 | 0.4599
|  | Randox Urine QC 3 | 1210UC | 34 |  |  | 807.118 | 0.585 | 0.55575 | 0.61425
|  | Ran-Liq 3 | 1238UC | 87 |  |  | 828.839 | 0.694 | 0.6593 | 0.7287
|  | Randox Urine QC 2 | 1209UC | 32 |  |  | 415.156 | 0.676 | 0.6422 | 0.7098
VITAMIN B12 II, U/L | C6000-A:E1-2 | [12 empty] | 6.188 | 5.8786 | 6.4974
|  | RIAP L2 | 2228EC | 91 |  |  | 443.052 | 3.455 | 3.2822 | 3.6277
|  | RIAP L1 | 2105EC | 436 |  |  | 171.022 | 7.437 | 7.0652 | 7.8089
|  | RIAP L2 | 2107EC | 407 |  |  | 441.334 | 4.036 | 3.8342 | 4.2378
|  | RIAP L1 | 2179EC | 190 |  |  | 159.453 | 6.255 | 5.9422 | 6.5678
|  | RIAP L2 | 2124EC | 115 |  |  | 436.328 | 8.434 | 8.0123 | 8.8557
proBNP, U/L | C6000-A:E1-1 | [12 empty] | 3.5861 | 3.4068 | 3.7654
|  | PC CARD1 | 657390 | 65 |  |  | 126.177 | 4.313 | 4.0973 | 4.5286
|  | PC CARD2 | 657389 | 63 |  |  | 4484.127 | 4.1 | 3.895 | 4.305
|  | PC CARD1 | 629283 | 63 |  |  | 128.321 | 3.826 | 3.6347 | 4.0173
|  | PC CARD1 | 586866 | 87 |  |  | 121.037 | 3.93 | 3.7335 | 4.1265
|  | PC CARD2 | 629284 | 61 |  |  | 4743.607 | 3.303 | 3.1378 | 3.4682
|  | PC CARD1 | 623415 | 163 |  |  | 123.507 | 3.259 | 3.096 | 3.4219
|  | PC CARD2 | 623416 | 162 |  |  | 4727.735 | 2.991 | 2.8415 | 3.1406
|  | PC CARD2 | 586867 | 87 |  |  | 4871.494 | 3.844 | 3.6518 | 4.0362
ADENOSINE DEAMINASE FLUID, U/L | C6000-B:C2 | [9 empty] | 4.2331 | 4.0214 | 4.4447 | 4.2331 | 4.0214 | 4.4447
|  | ADA 2 | 201220 | 132 |  |  | 142.541 | 4.45
|  | ADA 2 | ADC203210 | 46 |  |  | 140.348 | 4.632
|  | ADA 1 | ADC103210 | 117 |  |  | 31.75 | 4.901
|  | ADA 2 | 202210 | 102 |  |  | 155.545 | 3.846
|  | ADA 0 | 01220 | 131 |  |  | 11.458 | 5.337
|  | ADA 0 | ADCL02210 | 106 |  |  | 11.259 | 3.044
|  | ADA 1 | 101220 | 96 |  |  | 30.457 | 4.017
|  | ADA 1 | 102210 | 67 |  |  | 32.54 | 2.964
|  | ADA 0 | L01210 | 49 |  |  | 10.698 | 2.775
ADRENOCORTICOTROPIC HORMONE, U/L | C6000-A:E1-1 | [9 empty] | 4.9617 | 4.7136 | 5.2098 | 4.9617 | 4.7136 | 5.2098
|  | PC MM1 | 523425 | 40 |  |  | 8.821 | 5.165
|  | PC MM2 | 523426 | 39 |  |  | 194.382 | 4.744
ALANINE TRANSAMINASE, U/L | C6000-A:C1 | [9 empty] | 2.5754 | 2.4466 | 2.7042 | 3.8169 | 3.6261 | 4.0078
|  | PCCC 1 | 525027 | 329 |  |  | 45.889 | 2.932
|  | PCCC 2 | 535719 | 93 |  |  | 122.088 | 2.783
|  | s-PCCC 2 | 461597 | 39 |  |  | 121.946 | 2.104
|  | PCCC 2 | 519198 | 369 |  |  | 116.679 | 2.422
|  | s-PCCC 1 | 525027 | 47 |  |  | 47.455 | 2.757
|  | PCCC 1 | 525205 | 147 |  |  | 47.612 | 1.929
ALANINE TRANSAMINASE, U/L | C6000-B:C2 | [9 empty] | 4.7481 | 4.5107 | 4.9855 | 3.8169 | 3.6261 | 4.0078
|  | PCCC 1-B | 525205 | 158 |  |  | 47.963 | 2.775
|  | s-PCCC 2-B | 461597 | 41 |  |  | 118.668 | 1.212
|  | PCCC 2-B | 535719 | 98 |  |  | 122.306 | 10.557
|  | PCCC 2-B | 519198 | 365 |  |  | 116.165 | 3.943
|  | s-PCCC 1-B | 525027 | 46 |  |  | 46.1 | 1.53
|  | PCCC 1-B | 525027 | 311 |  |  | 45.36 | 4.014
ALBUMIN, U/L | C6000-A:C1 | [9 empty] | 2.5063 | 2.381 | 2.6316 | 2.8601 | 2.7171 | 3.0031
|  | PCCC 1 | 525205 | 141 |  |  | 32.523 | 3.197
|  | PCCC 2 | 519198 | 371 |  |  | 49.789 | 1.831
|  | PCCC 2 | 535719 | 88 |  |  | 48.601 | 2.165
|  | PCCC 2 | 461597 | 73 |  |  | 51.996 | 2.014
|  | PCCC 1 | 525027 | 407 |  |  | 31.602 | 2.886
ALBUMIN, U/L | C6000-B:C2 | [9 empty] | 3.077 | 2.9231 | 3.2308 | 2.8601 | 2.7171 | 3.0031
|  | s-PCCC 2-B | 519198 | 584 |  |  | 49.95 | 2.779
|  | s-PCCC 1-B | 525205 | 227 |  |  | 31.924 | 4.315
|  | s-PCCC 2-B | 535719 | 146 |  |  | 47.631 | 2.319
|  | PCCC 2-B | 461597 | 68 |  |  | 51.828 | 1.861
|  | s-PCCC 1-B | 525027 | 488 |  |  | 31.792 | 3.183
|  | PCCC 1-B | 525027 | 78 |  |  | 30.965 | 2.133
ALBUMIN CSF, U/L | C6000-A:C1 | [9 empty] | 3.8053 | 3.615 | 3.9955 | 3.8053 | 3.615 | 3.9955
|  | LIQ SFC2 | 56502 | 110 |  |  | 342.374 | 3.435
|  | LIQ SFC2 | 56492 | 48 |  |  | 350.658 | 1.765
|  | PPPUC TPC3 | 646927 | 121 |  |  | 121.158 | 3.996
|  | LIQ SFC2 | 56522 | 128 |  |  | 317.951 | 3.656
|  | PPPUC TPC3 | 637091 | 120 |  |  | 121.775 | 4.929
|  | PPPUC TPC3 | 531181 | 37 |  |  | 119.67 | 2.911
|  | LIQ SFC2 | 56532 | 39 |  |  | 306.392 | 2.941
|  | PPPUC TPC3 | 602710 | 157 |  |  | 127.754 | 4.332
|  | LIQ SFC2 | 56512 | 100 |  |  | 353.524 | 2.885
ALBUMIN URINE, U/L | C6000-A:C1 | [9 empty] | 10.5758 | 10.047 | 11.1046 | 10.5758 | 10.047 | 11.1046
|  | PN PUC | 602712 | 123 |  |  | 31.905 | 3.774
|  | PP PUC | 64692700 | 84 |  |  | 114.132 | 2.439
|  | PN PUC | 649811 | 183 |  |  | 31.864 | 18.891
|  | Ran-Liq 2 | 1237UC | 75 |  |  | 25.111 | 22.904
|  | PN PUC | 531179 | 82 |  |  | 27.938 | 3.603
|  | PP PUC | 531181 | 40 |  |  | 114.28 | 1.82
|  | Ran-Liq 2 | 1209UC | 68 |  |  | 23.709 | 9.198
|  | Ran-Liq 3 | 1238UC | 79 |  |  | 63.152 | 6.381
|  | Ran-Liq 3 | 1210UC | 68 |  |  | 63.488 | 3.215
|  | PP PUC | 637091 | 121 |  |  | 119.967 | 3.508
|  | PP PUC | 602710 | 161 |  |  | 122.05 | 3.37
ALDO, U/L | LIAISON XL | [9 empty] | 8.3729 | 7.9542 | 8.7915 | 8.3729 | 7.9542 | 8.7915
|  | ALDO L1 | 227012 | 51 |  |  | 189.294 | 10.004
|  | ALDO L2 | 228012 | 48 |  |  | 740.333 | 6.182
ALKALINE PHOSPHATASE, U/L | C6000-A:C1 | [9 empty] | 4.8878 | 4.6434 | 5.1322 | 4.8814 | 4.6374 | 5.1255
|  | s-PCCC 2 | 519198 | 619 |  |  | 239.817 | 4.231
|  | PCCC 2 | 461597 | 69 |  |  | 231.623 | 2.933
|  | s-PCCC 1 | 525205 | 268 |  |  | 100.791 | 5.523
|  | PCCC 1 | 525027 | 92 |  |  | 90.413 | 4.647
|  | s-PCCC 1 | 525027 | 520 |  |  | 91.083 | 5.308
|  | s-PCCC 2 | 535719 | 160 |  |  | 267.206 | 5.471
ALKALINE PHOSPHATASE, U/L | C6000-B:C2 | [9 empty] | 4.8747 | 4.631 | 5.1185 | 4.8814 | 4.6374 | 5.1255
|  | s-PCCC 2-B | 519198 | 588 |  |  | 237.776 | 4.367
|  | PCCC 2-B | 461597 | 68 |  |  | 230.235 | 2.774
|  | s-PCCC 1-B | 525205 | 254 |  |  | 97.413 | 6.573
|  | s-PCCC 1-B | 525027 | 484 |  |  | 90.426 | 4.822
|  | PCCC 1-B | 525027 | 80 |  |  | 88.175 | 5.11
|  | s-PCCC 2-B | 535719 | 162 |  |  | 258.068 | 4.17
ALPHA 1-ANTITRYPSIN, U/L | C6000-B:C2 | [9 empty] | 2.6384 | 2.5065 | 2.7703 | 2.6384 | 2.5065 | 2.7703
|  | PCCC 1-B | 525027 | 231 |  |  | 0.802 | 3.023
|  | PCCC 1-B | 525205 | 139 |  |  | 0.904 | 2.436
|  | s-PCCC 2-B | 461597 | 33 |  |  | 1.277 | 1.568
|  | s-PCCC 1-B | 525027 | 36 |  |  | 0.802 | 1.738
|  | PCCC 2-B | 519198 | 277 |  |  | 1.289 | 2.863
|  | PCCC 2-B | 535719 | 85 |  |  | 1.343 | 1.409
ALPHA-FETO PROTEIN, U/L | C6000-A:E1-1 | [9 empty] | 5.0977 | 4.8428 | 5.3525 | 5.0977 | 4.8428 | 5.3525
|  | RIAP L3 | 2147 | 158 |  |  | 231.45 | 6.611
|  | RIAP L3 | 2181EC | 145 |  |  | 205.785 | 4.099
|  | RIAP L1 | 2179EC | 139 |  |  | 10.187 | 5.934
|  | RIAP L1 | 2105EC | 311 |  |  | 11.715 | 4.779
|  | RIAP L3 | 2110EC | 149 |  |  | 231.58 | 3.757
AMMONIA, U/L | C6000-B:C2 | [9 empty] | 3.3831 | 3.214 | 3.5523 | 3.3831 | 3.214 | 3.5523
|  | AMM-P | 623128 | 81 |  |  | 233.749 | 2.111
|  | AMM-N | 644004 | 167 |  |  | 63.522 | 3.62
|  | AMM-N | 587880 | 232 |  |  | 63.501 | 4.143
|  | AMM-P | 587881 | 227 |  |  | 223.73 | 2.859
|  | AMM-P | 644006 | 87 |  |  | 227.843 | 2.825
ANTI-THYROGLOBULIN AB, U/L | C6000-A:E2-1 | [9 empty] | 6.6651 | 6.3318 | 6.9983 | 6.6651 | 6.3318 | 6.9983
|  | PC THYRO1 | 580444 | 228 |  |  | 112.67 | 7.186
|  | PC THYRO2 | 580446 | 231 |  |  | 221.702 | 7.653
|  | PC THYRO2 | 612140 | 110 |  |  | 217.754 | 5.703
|  | PC THYRO1 | 612139 | 87 |  |  | 101.275 | 3.899
|  | PC THYRO1 | 669238 | 49 |  |  | 62.587 | 4.679
ANTI-THYROID PEROXIDASE AB, U/L | C6000-A:E1-1 | [9 empty] | 8.848 | 8.4056 | 9.2904 | 8.848 | 8.4056 | 9.2904
|  | PC THYRO2 | 580446 | 42 |  |  | 109.415 | 7.653
|  | PC THYRO1 | 580444 | 43 |  |  | 47.726 | 9.876
APOPROTEIN A, U/L | C6000-B:C2 | [9 empty] | 4.5237 | 4.2975 | 4.7499 | 4.5237 | 4.2975 | 4.7499
|  | PCCC 1-B | 525027 | 98 |  |  | 1.119 | 4.366
|  | PCCC 2-B | 519198 | 116 |  |  | 1.747 | 4.51
|  | PCCC 1-B | 525205 | 38 |  |  | 1.153 | 4.952
APOPROTEIN B, U/L | C6000-B:C2 | [9 empty] | 5.0186 | 4.7676 | 5.2695 | 5.0186 | 4.7676 | 5.2695
|  | PCCC 1-B | 525205 | 39 |  |  | 0.475 | 5.424
|  | PCCC 2-B | 519198 | 110 |  |  | 0.764 | 4.809
|  | PCCC 1-B | 525027 | 93 |  |  | 0.471 | 5.087
ASPARTATE TRANSAMINASE, U/L | C6000-A:C1 | [9 empty] | 2.1164 | 2.0106 | 2.2223 | 3.6446 | 3.4624 | 3.8268
|  | PCCC 1 | 525205 | 141 |  |  | 44.284 | 1.641
|  | PCCC 2 | 535719 | 91 |  |  | 137.558 | 1.763
|  | s-PCCC 1 | 525027 | 43 |  |  | 46 | 2.012
|  | PCCC 2 | 519198 | 344 |  |  | 143.651 | 2.16
|  | s-PCCC 2 | 461597 | 37 |  |  | 148.027 | 1.37
|  | PCCC 1 | 525027 | 307 |  |  | 45.498 | 2.42
ASPARTATE TRANSAMINASE, U/L | C6000-B:C2 | [9 empty] | 4.8065 | 4.5662 | 5.0468 | 3.6446 | 3.4624 | 3.8268
|  | s-PCCC 2-B | 461597 | 33 |  |  | 149.979 | 2.409
|  | PCCC 1-B | 525205 | 139 |  |  | 44.905 | 2.804
|  | PCCC 2-B | 535719 | 88 |  |  | 139.222 | 11.021
|  | PCCC 2-B | 519198 | 302 |  |  | 146.775 | 3.509
|  | s-PCCC 1-B | 525027 | 38 |  |  | 46.361 | 2.259
|  | PCCC 1-B | 525027 | 264 |  |  | 46.514 | 3.943
BETA-HCG, U/L | C6000-A:E1-1 | [9 empty] | 4.7387 | 4.5018 | 4.9757 | 5.7593 | 5.4713 | 6.0473
|  | RIAP L3 | 2110EC | 151 |  |  | 300.732 | 4.254
|  | RIAP L3 | 2181EC | 151 |  |  | 252.719 | 5.082
|  | RIAP L1 | 2179EC | 147 |  |  | 11.284 | 5.528
|  | RIAP L1 | 2105EC | 324 |  |  | 11.02 | 4.668
|  | RIAP L3 | 2147 | 165 |  |  | 295.538 | 4.178
BETA-HCG, U/L | C6000-B:E3-2 | [9 empty] | 6.6788 | 6.3449 | 7.0127 | 5.7593 | 5.4713 | 6.0473
|  | RIAP L3-B | 2147EC | 148 |  |  | 305.552 | 6.47
|  | RIAP L3-B | 2181EC | 154 |  |  | 258.142 | 6.555
|  | RIAP L3-B | 2110EC | 144 |  |  | 294.176 | 4.566
|  | RIAP L1-B | 2105EC | 292 |  |  | 11.134 | 7.94
|  | RIAP L1-B | 2179EC | 141 |  |  | 11.468 | 5.935
BICARBONATE, U/L | C6000-B:C2 | [9 empty] | 4.4566 | 4.2338 | 4.6795 | 4.4566 | 4.2338 | 4.6795
|  | AMM-P | 644006 | 94 |  |  | 29.154 | 4.887
|  | AMM-P | 623128 | 65 |  |  | 29.68 | 3.804
|  | AMM-N | 587880 | 219 |  |  | 18.775 | 4.959
|  | AMM-P | 587881 | 214 |  |  | 32.164 | 3.387
|  | AMM-N | 644004 | 160 |  |  | 17.956 | 4.943
C-REACTIVE PROTEIN, U/L | C6000-A:C1 | [9 empty] | 3.129 | 2.9726 | 3.2855 | 4.7053 | 4.47 | 4.9405
|  | s-PCCC 1 | 525027 | 48 |  |  | 8.805 | 2.74
|  | PCCC 2 | 519198 | 366 |  |  | 50.944 | 3.259
|  | PCCC 1 | 525027 | 321 |  |  | 8.809 | 3.247
|  | s-PCCC 2 | 461597 | 43 |  |  | 52.091 | 4.701
|  | PCCC 2 | 535719 | 98 |  |  | 51.359 | 2.707
|  | PCCC 1 | 525205 | 156 |  |  | 5.663 | 2.276
C-REACTIVE PROTEIN, U/L | C6000-B:C2 | [9 empty] | 6.0542 | 5.7515 | 6.3569 | 4.7053 | 4.47 | 4.9405
|  | PCCC 2-B | 519198 | 316 |  |  | 51.332 | 3.251
|  | s-PCCC 1-B | 525027 | 41 |  |  | 8.731 | 2.074
|  | PCCC 1-B | 525027 | 261 |  |  | 8.765 | 4.046
|  | PCCC 1-B | 525205 | 139 |  |  | 5.53 | 9.803
|  | s-PCCC 2-B | 461597 | 37 |  |  | 54.266 | 2.549
|  | PCCC 2-B | 535719 | 85 |  |  | 50.955 | 11.31
C-peptide, U/L | C6000-A:E2-2 | [9 empty] | 4.7123 | 4.4767 | 4.9479 | 4.7123 | 4.4767 | 4.9479
|  | RIAP L2 | 2107EC | 357 |  |  | 6.266 | 3.51
|  | RIAP L1 | 2105EC | 382 |  |  | 2.103 | 4.315
|  | RIAP L1 | 2179EC | 185 |  |  | 1.776 | 7.279
|  | RIAP L2 | 2124EC | 105 |  |  | 6.684 | 4.689
|  | RIAP L2 | 2228EC | 89 |  |  | 6.068 | 3.692
CA 125, U/L | C6000-A:E2-1 | [9 empty] | 4.2448 | 4.0326 | 4.4571 | 4.2448 | 4.0326 | 4.4571
|  | RIAP L1 | 2179EC | 121 |  |  | 20.111 | 3.863
|  | RIAP L3 | 2147 | 155 |  |  | 190.874 | 5.527
|  | RIAP L1 | 2105EC | 302 |  |  | 18.505 | 4.625
|  | RIAP L3 | 2110EC | 139 |  |  | 187.761 | 3.253
|  | RIAP L3 | 2181EC | 133 |  |  | 195.694 | 2.51
CA 19-9, U/L | C6000-A:E2-1 | [9 empty] | 5.5818 | 5.3027 | 5.8609 | 5.5818 | 5.3027 | 5.8609
|  | RIAP L3 | 2110EC | 146 |  |  | 168.101 | 3.852
|  | RIAP L1 | 2179EC | 124 |  |  | 19.958 | 5.767
|  | RIAP L1 | 2105EC | 327 |  |  | 23.598 | 6.566
|  | RIAP L3 | 2181EC | 138 |  |  | 166.654 | 4.298
|  | RIAP L3 | 2147 | 164 |  |  | 162.957 | 5.553
CALCIUM, U/L | C6000-A:C1 | [9 empty] | 1.8144 | 1.7237 | 1.9051 | 2.9214 | 2.7753 | 3.0675
|  | PCCC 2 | 519198 | 378 |  |  | 3.395 | 1.949
|  | PCCC 2 | 535719 | 94 |  |  | 3.458 | 1.453
|  | PCCC 2 | 461597 | 70 |  |  | 3.391 | 0.904
|  | PCCC 1 | 525205 | 148 |  |  | 2.146 | 2.009
|  | PCCC 1 | 525027 | 409 |  |  | 2.237 | 1.801
CALCIUM, U/L | C6000-B:C2 | [9 empty] | 3.7176 | 3.5317 | 3.9035 | 2.9214 | 2.7753 | 3.0675
|  | PCCC 2-B | 519198 | 371 |  |  | 3.35 | 5.748
|  | PCCC 2-B | 535719 | 96 |  |  | 3.441 | 1.037
|  | PCCC 2-B | 461597 | 74 |  |  | 3.397 | 1.752
|  | PCCC 1-B | 525205 | 156 |  |  | 2.117 | 2.176
|  | PCCC 1-B | 525027 | 393 |  |  | 2.22 | 2.089
CALCIUM URINE, U/L | C6000-B:C2 | [9 empty] | 2.4815 | 2.3575 | 2.6056 | 2.4815 | 2.3575 | 2.6056
|  | Ran-Liq 3 | 1238UC | 127 |  |  | 2.859 | 2.866
|  | Ran-Liq 2 | 1209UC | 224 |  |  | 1.979 | 2.219
|  | Ran-Liq 3 | 1210UC | 224 |  |  | 2.897 | 2.505
|  | Bio-liq1 | 88161 | 122 |  |  | 1.826 | 2.447
|  | Bio-liq 2 | 88162 | 112 |  |  | 2.71 | 1.74
|  | Ran-Liq 2 | 1237UC | 116 |  |  | 1.925 | 3.045
CARCINOEMBRYONIC AG, U/L | C6000-A:E1-2 | [9 empty] | 5.6745 | 5.3908 | 5.9583 | 5.6745 | 5.3908 | 5.9583
|  | RIAP L1 | 2105EC | 314 |  |  | 4.379 | 3.964
|  | RIAP L1 | 2179EC | 120 |  |  | 4.204 | 6.127
|  | RIAP L3 | 2181EC | 134 |  |  | 50.868 | 8.219
|  | RIAP L3 | 2147 | 151 |  |  | 52.876 | 6.699
|  | RIAP L3 | 2110EC | 143 |  |  | 53.743 | 4.191
CHLORIDE, U/L | C6000-A:ISE1 | [9 empty] | 1.9452 | 1.8479 | 2.0424 | 3.3173 | 3.1515 | 3.4832
|  | PCCC 1 | 525027 | 106 |  |  | 87.111 | 1.611
|  | PCCC 2 | 461597 | 88 |  |  | 103.232 | 1.747
|  | s-PCCC 1 | 525205 | 342 |  |  | 84.446 | 2.174
|  | s-PCCC 2 | 519198 | 712 |  |  | 105.235 | 1.747
|  | s-PCCC 1 | 525027 | 595 |  |  | 87.555 | 2.184
|  | s-PCCC 2 | 535719 | 193 |  |  | 105.448 | 1.658
CHLORIDE, U/L | C6000-B:ISE1 | [9 empty] | 4.112 | 3.9064 | 4.3176 | 3.3173 | 3.1515 | 3.4832
|  | PCCC 1-B | 525027 | 368 |  |  | 88.191 | 1.762
|  | PCCC 2-B | 519198 | 278 |  |  | 105.84 | 1.499
|  | PCCC 2-B | 461597 | 92 |  |  | 103.572 | 1.328
|  | s-PCCC 2-B | 535719 | 188 |  |  | 106.224 | 3.356
|  | s-PCCC 1-B | 525027 | 565 |  |  | 88.418 | 1.773
|  | s-PCCC 1-B | 525205 | 303 |  |  | 85.98 | 10.585
|  | s-PCCC 2-B | 519198 | 696 |  |  | 106.061 | 1.837
CHLORIDE URINE, U/L | C6000-B:ISE1 | [9 empty] | 3.2449 | 3.0826 | 3.4071 | 3.2449 | 3.0826 | 3.4071
|  | Bio-liq1 | 88161 | 165 |  |  | 93.919 | 2.229
|  | Bio-liq 2 | 88162 | 157 |  |  | 194.188 | 2.051
|  | Ran-Liq 3 | 1238UC | 154 |  |  | 180.447 | 2.614
|  | Ran-Liq 3 | 1210UC | 252 |  |  | 168.146 | 3.876
|  | Ran-Liq 2 | 1237UC | 138 |  |  | 79.582 | 2.111
|  | Ran-Liq 2 | 1209UC | 246 |  |  | 68.199 | 4.402
COMPLEMENT C3, U/L | C6000-B:C2 | [9 empty] | 3.7153 | 3.5295 | 3.9011 | 3.7153 | 3.5295 | 3.9011
|  | PCCC 2-B | 519198 | 248 |  |  | 1.588 | 3.932
|  | s-PCCC 1-B | 525027 | 30 |  |  | 0.961 | 1.612
|  | PCCC 1-B | 525027 | 212 |  |  | 0.974 | 3.035
|  | PCCC 1-B | 525205 | 109 |  |  | 0.938 | 4.16
|  | PCCC 2-B | 535719 | 65 |  |  | 1.507 | 4.66
COMPLEMENT C4, U/L | C6000-B:C2 | [9 empty] | 2.969 | 2.8205 | 3.1174 | 2.969 | 2.8205 | 3.1174
|  | PCCC 1-B | 525205 | 109 |  |  | 0.139 | 2.196
|  | PCCC 1-B | 525027 | 195 |  |  | 0.124 | 3.519
|  | PCCC 2-B | 519198 | 231 |  |  | 0.218 | 3.09
|  | PCCC 2-B | 535719 | 66 |  |  | 0.223 | 1.484
CONJUGATED BILIRUBIN, U/L | C6000-A:C1 | [9 empty] | 4.763 | 4.5248 | 5.0011 | 4.5514 | 4.3239 | 4.779
|  | PCCC 2 | 461597 | 69 |  |  | 41.146 | 4.491
|  | PCCC 1 | 525027 | 393 |  |  | 14.943 | 6.381
|  | MAS BILIRUBIN 3 | BC24033A | 386 |  |  | 172.779 | 4.039
|  | PCCC 1 | 525205 | 151 |  |  | 16.193 | 3.551
|  | PCCC 2 | 519198 | 354 |  |  | 41.764 | 4.276
|  | PCCC 2 | 535719 | 98 |  |  | 44.05 | 2.994
|  | MAS BILIRUBIN 3 | BC23063A | 66 |  |  | 167.064 | 4.686
CONJUGATED BILIRUBIN, U/L | C6000-B:C2 | [9 empty] | 4.3084 | 4.093 | 4.5238 | 4.5514 | 4.3239 | 4.779
|  | PCCC 1-B | 525027 | 354 |  |  | 15.474 | 4.739
|  | PCCC 2-B | 535719 | 93 |  |  | 43.499 | 3.587
|  | PCCC 2-B | 519198 | 328 |  |  | 42.629 | 4.937
|  | PCCC 1-B | 525205 | 148 |  |  | 16.161 | 3.781
|  | MAS BILIRUBIN-B | BC24033A | 338 |  |  | 174.774 | 4.088
|  | PCCC 2-B | 461597 | 67 |  |  | 41.71 | 2.253
|  | MAS BILIRUBIN-B | BC23063A | 61 |  |  | 170.185 | 2.896
CORTISOL II, U/L | C6000-A:E1-2 | [9 empty] | 4.6383 | 4.4064 | 4.8702 | 4.6383 | 4.4064 | 4.8702
|  | RIAP L1 | 2105EC | 300 |  |  | 197.438 | 4.144
|  | RIAP L2 | 2228EC | 63 |  |  | 694.73 | 6.147
|  | RIAP L1 | 2179EC | 117 |  |  | 198.93 | 5.838
|  | RIAP L2 | 2124EC | 74 |  |  | 641.069 | 5.738
|  | RIAP L2 | 2107EC | 281 |  |  | 643.459 | 3.79
CREATINE KINASE Gen 2, U/L | C6000-A:C1 | [9 empty] | 1.8516 | 1.759 | 1.9442 | 1.8516 | 1.759 | 1.9442
|  | s-PCCC 2 | 461597 | 36 |  |  | 259.417 | 1.368
|  | PCCC 1 | 525027 | 289 |  |  | 149.512 | 2.159
|  | s-PCCC 1 | 525027 | 44 |  |  | 149.636 | 1.117
|  | PCCC 1 | 525205 | 134 |  |  | 148.224 | 1.401
|  | PCCC 2 | 535719 | 82 |  |  | 261.11 | 1.443
|  | PCCC 2 | 519198 | 329 |  |  | 256.502 | 1.93
CREATININE, U/L | C6000-A:C1 | [9 empty] | 2.0919 | 1.9873 | 2.1965 | 2.14 | 2.033 | 2.247
|  | PCCC 2 | 461597 | 136 |  |  | 361.184 | 1.648
|  | PCCC 1 | 525205 | 417 |  |  | 89.326 | 2.232
|  | PCCC 1 | 525027 | 979 |  |  | 89.396 | 2.292
|  | PCCC 2 | 535719 | 259 |  |  | 330.95 | 2.131
|  | PCCC 2 | 519198 | 982 |  |  | 352.287 | 1.852
CREATININE, U/L | C6000-B:C2 | [9 empty] | 2.1984 | 2.0884 | 2.3083 | 2.14 | 2.033 | 2.247
|  | PCCC 1-B | 525027 | 762 |  |  | 88.979 | 2.387
|  | PCCC 2-B | 519198 | 759 |  |  | 347.335 | 2.265
|  | PCCC 2-B | 461597 | 121 |  |  | 351.909 | 2.525
|  | PCCC 1-B | 525205 | 359 |  |  | 87.518 | 1.851
|  | PCCC 2-B | 535719 | 231 |  |  | 322.996 | 1.541
CREATININE URINE, U/L | C6000-A:C1 | [9 empty] | 2.7822 | 2.643 | 2.9213 | 2.7193 | 2.5833 | 2.8552
|  | Ran-Liq 3 | 1210UC | 119 |  |  | 14.508 | 3.757
|  | Ran-Liq 2 | 1237UC | 114 |  |  | 6.48 | 2.953
|  | PP PUC | 64692700 | 186 |  |  | 4.558 | 2.128
|  | PN PUC | 649811 | 401 |  |  | 8.843 | 2.072
|  | Ran-Liq 2 | 1209UC | 124 |  |  | 6.357 | 3.352
|  | PN PUC | 602712 | 295 |  |  | 8.525 | 2.746
|  | PN PUC | 531179 | 174 |  |  | 8.619 | 3.803
|  | PN PUC | 634291 | 37 |  |  | 8.92 | 1.365
|  | PP PUC | 602710 | 385 |  |  | 4.531 | 2.867
|  | PP PUC | 531181 | 80 |  |  | 4.451 | 3.702
|  | Ran-Liq 3 | 1238UC | 112 |  |  | 14.759 | 2.615
|  | PP PUC | 637091 | 239 |  |  | 4.522 | 2.125
CREATININE URINE, U/L | C6000-B:C2 | [9 empty] | 2.3977 | 2.2778 | 2.5176 | 2.7193 | 2.5833 | 2.8552
|  | Ran-Liq 3 | 1238UC | 33 |  |  | 15.144 | 1.196
|  | Ran-Liq 2 | 1237UC | 145 |  |  | 6.536 | 2.126
|  | Ran-Liq 2 | 1209UC | 299 |  |  | 6.431 | 2.609
DEHYDROEPIANDROSTERONE SULPHAT, U/L | C6000-A:E2-2 | [9 empty] | 5.5647 | 5.2864 | 5.8429 | 5.5647 | 5.2864 | 5.8429
|  | RIAP L2 | 2107EC | 277 |  |  | 15.98 | 4.841
|  | RIAP L1 | 2179EC | 117 |  |  | 2.635 | 6.794
|  | RIAP L2 | 2124EC | 80 |  |  | 16.85 | 4.971
|  | RIAP L1 | 2105EC | 291 |  |  | 2.795 | 5.394
|  | RIAP L2 | 2228EC | 68 |  |  | 17.321 | 7.152
FERRITIN, U/L | C6000-A:E2-1 | [9 empty] | 6.1978 | 5.8879 | 6.5077 | 6.1978 | 5.8879 | 6.5077
|  | RIAP L3 | 2181EC | 163 |  |  | 340.861 | 4.244
|  | sRIAP L1 | 2179EC | 290 |  |  | 22.798 | 6.977
|  | PC V1 | 595763 | 44 |  |  | 144.439 | 2.045
|  | sRIAP L3 | 2147 EC | 325 |  |  | 311.743 | 3.752
|  | sRIAP L3 | 2181EC | 307 |  |  | 339.172 | 3.731
|  | sRIAP L1 | 2105EC | 515 |  |  | 14.33 | 8.887
|  | RIAP L1 | 2105EC | 133 |  |  | 15.711 | 5.935
|  | RIAP L3 | 2110EC | 133 |  |  | 326.623 | 3.723
FOLLICLE STIMULATING HORMONE, U/L | C6000-A:E1-1 | [9 empty] | 3.3648 | 3.1965 | 3.533 | 3.2759 | 3.1121 | 3.4396
|  | RIAP L1 | 2105EC | 297 |  |  | 6.689 | 3.225
|  | RIAP L1 | 2179EC | 68 |  |  | 5.974 | 3.474
|  | RIAP L3 | 2147 | 148 |  |  | 54.04 | 2.743
|  | RIAP L3 | 2181EC | 71 |  |  | 52.477 | 2.777
|  | RIAP L3 | 2110EC | 140 |  |  | 56.794 | 4.331
FOLLICLE STIMULATING HORMONE, U/L | C6000-A:E1-2 | [9 empty] | 2.6874 | 2.5531 | 2.8218 | 3.2759 | 3.1121 | 3.4396
|  | RIAP L3 | 2181EC | 61 |  |  | 56.469 | 2.516
|  | RIAP L1 | 2179EC | 62 |  |  | 6.502 | 2.846
FREE PSA, U/L | C6000-A:E1-2 | [9 empty] | 4.891 | 4.6465 | 5.1356 | 4.891 | 4.6465 | 5.1356
|  | RIAP L1 | 2179EC | 114 |  |  | 1.154 | 3.884
|  | RIAP L2 | 2228EC | 60 |  |  | 9.151 | 2.554
|  | RIAP L2 | 2107EC | 290 |  |  | 7.716 | 5.567
|  | RIAP L2 | 2124EC | 75 |  |  | 10.005 | 2.861
|  | RIAP L1 | 2105EC | 302 |  |  | 1.225 | 5.266
Free T4 Gen 3, U/L | C6000-A:E1-1 | [9 empty] | 3.8991 | 3.7041 | 4.094 | 3.8991 | 3.7041 | 4.094
|  | sRIAP L2 | 2228EC | 126 |  |  | 29.033 | 4.191
|  | RIAP L2 | 2107EC | 144 |  |  | 28.582 | 4.649
|  | sRIAP L1 | 2105EC | 476 |  |  | 13.497 | 3.737
|  | sRIAP L2 | 2107EC | 434 |  |  | 29.116 | 4.008
|  | sRIAP L2 | 2124EC | 147 |  |  | 29.66 | 2.48
|  | sRIAP L1 | 2179EC | 250 |  |  | 13.597 | 4.209
|  | RIAP L1 | 2105EC | 131 |  |  | 13.084 | 3.59
GAMMA-GLUTAMYL TRANSFERASE, U/L | C6000-A:C1 | [9 empty] | 2.0751 | 1.9714 | 2.1789 | 2.3619 | 2.2438 | 2.48
|  | PCCC 2 | 519198 | 349 |  |  | 220.934 | 1.616
|  | s-PCCC 1 | 525027 | 52 |  |  | 55.212 | 3.447
|  | s-PCCC 2 | 461597 | 45 |  |  | 271.067 | 3.772
|  | PCCC 1 | 525027 | 311 |  |  | 55.1 | 2.087
|  | PCCC 1 | 525205 | 142 |  |  | 57.408 | 2.087
|  | PCCC 2 | 535719 | 90 |  |  | 226.7 | 1.198
GAMMA-GLUTAMYL TRANSFERASE, U/L | C6000-B:C2 | [9 empty] | 2.6283 | 2.4969 | 2.7598 | 2.3619 | 2.2438 | 2.48
|  | s-PCCC 2-B | 461597 | 40 |  |  | 274.125 | 0.819
|  | PCCC 2-B | 519198 | 337 |  |  | 221.421 | 2.999
|  | PCCC 1-B | 525205 | 144 |  |  | 57.014 | 2.095
|  | s-PCCC 1-B | 525027 | 47 |  |  | 55.532 | 1.397
|  | PCCC 1-B | 525027 | 291 |  |  | 54.784 | 2.958
|  | PCCC 2-B | 535719 | 88 |  |  | 224.591 | 1.507
GLUCOSE, U/L | C6000-A:C1 | [9 empty] | 2.0622 | 1.9591 | 2.1653 | 2.2216 | 2.1105 | 2.3327
|  | PCCC 1 | 525205 | 130 |  |  | 5.787 | 2.197
|  | PCCC 2 | 535719 | 82 |  |  | 13.45 | 2.27
|  | PCCC 2 | 519198 | 326 |  |  | 13.47 | 2.056
|  | PCCC 2 | 461597 | 66 |  |  | 13.1 | 1.562
|  | PCCC 1 | 525027 | 360 |  |  | 5.759 | 2.048
GLUCOSE, U/L | C6000-B:C2 | [9 empty] | 2.3768 | 2.2579 | 2.4956 | 2.2216 | 2.1105 | 2.3327
|  | PCCC 1-B | 525205 | 132 |  |  | 5.769 | 2.927
|  | PCCC 2-B | 519198 | 311 |  |  | 13.393 | 2.562
|  | PCCC 2-B | 461597 | 62 |  |  | 12.974 | 1.572
|  | PCCC 2-B | 535719 | 81 |  |  | 13.277 | 2.14
|  | PCCC 1-B | 525027 | 337 |  |  | 5.701 | 2.121
GLYCATED HAEMOGLOBIN (NGSP), U/L | C6000-A:C1 | [9 empty] | 2.2841 | 2.1699 | 2.3983 | 1.7988 | 1.7088 | 1.8887
|  | PCA1N-VER | 617428 | 139 |  |  | 5.737 | 2.059
|  | PCA1P | 617429 | 91 |  |  | 9.646 | 1.709
|  | PCA1N-A | 563450 | 190 |  |  | 5.794 | 2.552
|  | PCA1P-A | 563451 | 191 |  |  | 10.587 | 2.393
GLYCATED HAEMOGLOBIN (NGSP), U/L | C6000-B:C2 | [9 empty] | 1.6121 | 1.5315 | 1.6927 | 1.7988 | 1.7088 | 1.8887
|  | PCA1N | 563450 | 63 |  |  | 5.737 | 1.213
|  | PCA1P | 563451 | 589 |  |  | 10.703 | 1.508
|  | PCA1P | 617429 | 245 |  |  | 9.737 | 1.76
|  | PCA1N | 509111 | 110 |  |  | 5.76 | 1.29
|  | PCA1N | 617428 | 357 |  |  | 5.761 | 1.891
|  | PCA1N | 548925 | 422 |  |  | 5.731 | 1.64
|  | PCA1P | 509113 | 110 |  |  | 10.86 | 1.126
HAPTOGLOBIN, U/L | C6000-B:C2 | [9 empty] | 2.5739 | 2.4452 | 2.7026 | 2.5739 | 2.4452 | 2.7026
|  | s-PCCC 2-B | 461597 | 36 |  |  | 1.308 | 1.381
|  | s-PCCC 1-B | 525027 | 43 |  |  | 0.765 | 1.216
|  | PCCC 2-B | 519198 | 332 |  |  | 1.221 | 2.94
|  | PCCC 1-B | 525205 | 156 |  |  | 0.8 | 2.256
|  | PCCC 2-B | 535719 | 104 |  |  | 1.253 | 1.312
|  | PCCC 1-B | 525027 | 284 |  |  | 0.782 | 2.87
HDL CHOLESTEROL, U/L | C6000-A:C1 | [9 empty] | 3.4043 | 3.2341 | 3.5745 | 3.4043 | 3.2341 | 3.5745
|  | PCCC 1 | 525027 | 283 |  |  | 0.764 | 3.124
|  | PCCC 2 | 519198 | 316 |  |  | 1.457 | 2.962
|  | PCCC 2 | 535719 | 84 |  |  | 1.595 | 4.044
|  | s-PCCC 2 | 461597 | 32 |  |  | 1.431 | 1.527
|  | PCCC 1 | 525205 | 131 |  |  | 0.753 | 4.857
|  | s-PCCC 1 | 525027 | 39 |  |  | 0.784 | 2.145
HIGH-SENSITIVITY TROPONIN T, U/L | C6000-A:E2-1 | [9 empty] | 5.965 | 5.6667 | 6.2632 | 6.0199 | 5.7189 | 6.3209
|  | Ran Trop T | 4607CK | 198 |  |  | 6.474 | 10.625
|  | PC TN1 | 590216 | 163 |  |  | 27.711 | 3.068
|  | BIO CTC1 | 89041 | 160 |  |  | 18.187 | 3.927
|  | PC TN1 | 558106 | 82 |  |  | 25.588 | 2.724
|  | PC TN1 | 629183 | 164 |  |  | 28.741 | 3.304
|  | BIO CTC1 | 89021 | 53 |  |  | 18.186 | 2.653
HIGH-SENSITIVITY TROPONIN T, U/L | C6000-B:E3-1 | [9 empty] | 6.0728 | 5.7691 | 6.3764 | 6.0199 | 5.7189 | 6.3209
|  | PC TN1-B | 590216 | 157 |  |  | 27.803 | 5.887
|  | PC TN1-B | 558106 | 86 |  |  | 24.246 | 2.047
|  | BIO CTC1-B | 89041 | 181 |  |  | 18.932 | 4.309
|  | Ran Trop T-B | 4607CK | 197 |  |  | 6.285 | 9.793
|  | PC TN1-B | 629183 | 176 |  |  | 29.951 | 4.046
|  | BIO CTC1-B | 89021 | 49 |  |  | 17.173 | 2.098
Human growth hormone, U/L | LIAISON XL | [9 empty] | 3.7499 | 3.5624 | 3.9374 | 3.7499 | 3.5624 | 3.9374
|  | hGH L2 | 9260310 | 52 |  |  | 17.45 | 3.568
|  | hGH L1 | 7980310 | 55 |  |  | 3.771 | 3.914
IGG CSF, U/L | C6000-A:C1 | [9 empty] | 4.2849 | 4.0707 | 4.4992 | 4.2849 | 4.0707 | 4.4992
|  | PNPUC TPC3 | 649811 | 157 |  |  | 21.281 | 4.28
|  | PNPUC TPC3 | 531179 | 75 |  |  | 20.031 | 5.101
|  | PPPUC TPC3 | 531181 | 32 |  |  | 134.157 | 3.374
|  | PNPUC TPC3 | 602712 | 109 |  |  | 23.003 | 6.289
|  | PNPUC TPC3 | 673292 | 35 |  |  | 19.981 | 4.753
|  | PPPUC TPC3 | 602710 | 154 |  |  | 133.509 | 3.278
|  | PPPUC TPC3 | 637091 | 101 |  |  | 142.362 | 2.534
|  | PPPUC TPC3 | 646927 | 112 |  |  | 128.098 | 3.821
IMMUNOGLOBULIN A, U/L | C6000-B:C2 | [9 empty] | 2.5071 | 2.3817 | 2.6324 | 2.5071 | 2.3817 | 2.6324
|  | PCCC 2-B | 535719 | 63 |  |  | 2.234 | 1.252
|  | PCCC 1-B | 525027 | 195 |  |  | 1.538 | 2.907
|  | PCCC 1-B | 525205 | 103 |  |  | 1.498 | 2.443
|  | PCCC 2-B | 519198 | 228 |  |  | 2.458 | 2.422
IMMUNOGLOBULIN G, U/L | C6000-B:C2 | [9 empty] | 2.7191 | 2.5832 | 2.8551 | 2.7191 | 2.5832 | 2.8551
|  | PCCC 2-B | 519198 | 226 |  |  | 11.94 | 2.662
|  | PCCC 1-B | 525205 | 108 |  |  | 7.688 | 2.515
|  | PCCC 2-B | 535719 | 68 |  |  | 11.649 | 2.352
|  | PCCC 1-B | 525027 | 193 |  |  | 7.382 | 2.878
|  | s-PCCC 1-B | 525027 | 32 |  |  | 7.437 | 3.439
IMMUNOGLOBULIN M, U/L | C6000-B:C2 | [9 empty] | 4.0695 | 3.866 | 4.273 | 4.0695 | 3.866 | 4.273
|  | PCCC 1-B | 525205 | 109 |  |  | 0.673 | 4.026
|  | PCCC 2-B | 535719 | 67 |  |  | 1.059 | 2.29
|  | PCCC 2-B | 519198 | 236 |  |  | 1.09 | 3.663
|  | PCCC 1-B | 525027 | 199 |  |  | 0.663 | 4.925
INORGANIC PHOSPHATE URINE, U/L | C6000-B:C2 | [9 empty] | 2.2636 | 2.1504 | 2.3768 | 2.2636 | 2.1504 | 2.3768
|  | Ran-Liq 2 | 1209UC | 213 |  |  | 8.734 | 2.433
|  | Bio-liq1 | 88161 | 113 |  |  | 8.302 | 2.051
|  | Bio-liq 2 | 88162 | 107 |  |  | 16.425 | 2.364
|  | Ran-Liq 3 | 1238UC | 127 |  |  | 14.63 | 2.297
|  | Ran-Liq 3 | 1210UC | 214 |  |  | 14.385 | 2.232
|  | Ran-Liq 2 | 1237UC | 115 |  |  | 8.722 | 2.052
INSULIN, U/L | C6000-A:E1-2 | [9 empty] | 5.7354 | 5.4486 | 6.0222 | 5.7354 | 5.4486 | 6.0222
|  | RIAP L3 | 2110EC | 156 |  |  | 27.819 | 5.737
|  | RIAP L3 | 2147 | 164 |  |  | 26.598 | 2.509
|  | RIAP L3 | 2181EC | 165 |  |  | 28.213 | 4.514
|  | RIAP L2 | 2228EC | 81 |  |  | 15.966 | 7.071
|  | RIAP L2 | 2124EC | 98 |  |  | 11.855 | 10.657
|  | RIAP L2 | 2107EC | 316 |  |  | 13.926 | 4.897
IRON, U/L | C6000-A:C1 | [9 empty] | 2.5802 | 2.4512 | 2.7092 | 2.5802 | 2.4512 | 2.7092
|  | PCCC 2 | 461597 | 67 |  |  | 43.791 | 1.625
|  | PCCC 2 | 535719 | 93 |  |  | 43.69 | 1.794
|  | PCCC 2 | 519198 | 375 |  |  | 43.156 | 2.031
|  | PCCC 1 | 525205 | 152 |  |  | 19.509 | 2.404
|  | PCCC 1 | 525027 | 401 |  |  | 19.892 | 3.288
Insulin-like growth factor I, U/L | LIAISON XL | [9 empty] | 15.8222 | 15.0311 | 16.6134 | 15.8222 | 15.0311 | 16.6134
|  | IGF-1 L2 | 1422081 | 36 |  |  | 233.297 | 17.711
|  | IGF-1 L2 | 1422084 | 37 |  |  | 238.219 | 15.914
|  | IGF-1 L1 | 1412084 | 35 |  |  | 60.231 | 13.493
LACTATE, U/L | C6000-B:C2 | [9 empty] | 2.6957 | 2.5609 | 2.8305 | 2.6957 | 2.5609 | 2.8305
|  | PCCC 1-B | 525027 | 231 |  |  | 1.76 | 2.952
|  | PCCC 2-B | 535719 | 83 |  |  | 3.811 | 1.474
|  | s-PCCC 1-B | 525027 | 33 |  |  | 1.754 | 1.548
|  | PCCC 1-B | 525205 | 135 |  |  | 1.781 | 3.131
|  | s-PCCC 2-B | 461597 | 30 |  |  | 3.658 | 0.942
|  | PCCC 2-B | 519198 | 271 |  |  | 3.818 | 2.754
LACTATE DEHYDROGENASE, U/L | C6000-A:C1 | [9 empty] | 2.445 | 2.3228 | 2.5673 | 2.6682 | 2.5347 | 2.8016
|  | s-PCCC 1 | 525027 | 47 |  |  | 175.106 | 1.137
|  | PCCC 2 | 519198 | 361 |  |  | 308.645 | 2.192
|  | PCCC 2 | 535719 | 98 |  |  | 301.398 | 2.295
|  | s-PCCC 2 | 461597 | 40 |  |  | 310.375 | 1.036
|  | PCCC 1 | 525027 | 321 |  |  | 173.555 | 3.006
|  | PCCC 1 | 525205 | 155 |  |  | 165.316 | 2.336
LACTATE DEHYDROGENASE, U/L | C6000-B:C2 | [9 empty] | 2.9142 | 2.7685 | 3.0599 | 2.6682 | 2.5347 | 2.8016
|  | PCCC 2-B | 535719 | 88 |  |  | 297.455 | 2.529
|  | PCCC 2-B | 519198 | 295 |  |  | 312.736 | 3.101
|  | PCCC 1-B | 525027 | 256 |  |  | 175.336 | 2.555
|  | PCCC 1-B | 525205 | 138 |  |  | 166.159 | 3.786
|  | s-PCCC 1-B | 525027 | 39 |  |  | 174.769 | 1.528
|  | s-PCCC 2-B | 461597 | 34 |  |  | 308.882 | 1.423
LIPASE, U/L | C6000-A:C1 | [9 empty] | 3.3757 | 3.2069 | 3.5445 | 3.3757 | 3.2069 | 3.5445
|  | s-PCCC 2 | 461597 | 43 |  |  | 93.747 | 3.123
|  | PCCC 2 | 535719 | 100 |  |  | 97.478 | 3.128
|  | PCCC 1 | 525027 | 338 |  |  | 44.378 | 3.547
|  | s-PCCC 1 | 525027 | 47 |  |  | 43.721 | 2.308
|  | PCCC 2 | 519198 | 382 |  |  | 100.585 | 3.558
|  | PCCC 1 | 525205 | 161 |  |  | 44.065 | 3.01
LIPOPROTEIN A2, U/L | C6000-B:C2 | [9 empty] | 5.5129 | 5.2372 | 5.7885 | 5.5129 | 5.2372 | 5.7885
|  | PC LPA 1 | 520719 | 95 |  |  | 42.46 | 4.809
|  | PC LPA 2 | 520720 | 92 |  |  | 108.754 | 6.156
LITHIUM, U/L | C6000-B:C2 | [9 empty] | 3.769 | 3.5806 | 3.9575 | 3.769 | 3.5806 | 3.9575
|  | s-PCCC 1-B | 525027 | 37 |  |  | 0.939 | 2.613
|  | PCCC 2-B | 519198 | 276 |  |  | 1.794 | 3.096
|  | PCCC 1-B | 525205 | 135 |  |  | 0.919 | 4.362
|  | s-PCCC 2-B | 461597 | 34 |  |  | 1.796 | 2.554
|  | PCCC 2-B | 535719 | 85 |  |  | 1.795 | 3.705
|  | PCCC 1-B | 525027 | 243 |  |  | 0.908 | 4.37
LUTEINISING HORMONE, U/L | C6000-A:E1-2 | [9 empty] | 4.2035 | 3.9933 | 4.4137 | 4.2035 | 3.9933 | 4.4137
|  | RIAP L3 | 2181EC | 132 |  |  | 54.598 | 4.678
|  | RIAP L3 | 2110EC | 136 |  |  | 53.477 | 3.993
|  | RIAP L1 | 2105EC | 298 |  |  | 3.576 | 3.565
|  | RIAP L1 | 2179EC | 125 |  |  | 3.748 | 5.852
|  | RIAP L3 | 2147 | 154 |  |  | 55.878 | 3.432
MAGNESIUM, U/L | C6000-A:C1 | [9 empty] | 2.2016 | 2.0915 | 2.3117 | 2.5527 | 2.4251 | 2.6804
|  | PCCC 2 | 461597 | 71 |  |  | 1.283 | 1.348
|  | PCCC 1 | 525027 | 417 |  |  | 0.792 | 2.373
|  | PCCC 1 | 525205 | 165 |  |  | 0.79 | 2.317
|  | PCCC 2 | 535719 | 103 |  |  | 1.348 | 2.167
|  | PCCC 2 | 519198 | 389 |  |  | 1.35 | 2.091
MAGNESIUM, U/L | C6000-B:C2 | [9 empty] | 2.8742 | 2.7305 | 3.0179 | 2.5527 | 2.4251 | 2.6804
|  | PCCC 2-B | 535719 | 99 |  |  | 1.346 | 3.716
|  | PCCC 1-B | 525205 | 168 |  |  | 0.79 | 3.907
|  | PCCC 2-B | 461597 | 71 |  |  | 1.294 | 1.677
|  | PCCC 1-B | 525027 | 388 |  |  | 0.794 | 2.725
|  | PCCC 2-B | 519198 | 370 |  |  | 1.355 | 2.351
MAGNESIUM URINE, U/L | C6000-B:C2 | [9 empty] | 7.1869 | 6.8276 | 7.5463 | 7.1869 | 6.8276 | 7.5463
|  | Ran-Liq 2 | 1209UC | 215 |  |  | 2.63 | 11.858
|  | Ran-Liq 3 | 1210UC | 216 |  |  | 5.35 | 6.355
|  | Ran-Liq 3 | 1238UC | 134 |  |  | 6.415 | 3.342
|  | Bio-liq 2 | 88162 | 125 |  |  | 3.925 | 2.894
|  | Ran-Liq 2 | 1237UC | 122 |  |  | 3.091 | 3.845
|  | Bio-liq1 | 88161 | 126 |  |  | 1.652 | 6.358
OESTRADIOL III, U/L | C6000-A:E1-1 | [9 empty] | 7.5589 | 7.181 | 7.9369 | 7.5589 | 7.181 | 7.9369
|  | RIAP L1 | 2179EC | 130 |  |  | 201.844 | 7.724
|  | RIAP L3 | 2110EC | 143 |  |  | 1245.881 | 3.8
|  | RIAP L3 | 2147 | 157 |  |  | 1507.191 | 3.164
|  | RIAP L3 | 2181EC | 133 |  |  | 1512.406 | 2.364
|  | RIAP L1 | 2105EC | 303 |  |  | 176.634 | 11.091
Osmolality, U/L | OSMO1 | [9 empty] | 0.90118 | 0.85612 | 0.94624 | 0.90118 | 0.85612 | 0.94624
|  | Serim BIOLIQ 1 | 45871 | 240 |  |  | 305.983 | 0.908
|  | Serim BIOLIQ 1 | 45931 | 114 |  |  | 311.789 | 0.898
|  | Serim BIOLIQ 2 | 45952 | 74 |  |  | 400.23 | 1.146
|  | Serim BIOLIQ 2 | 45872 | 110 |  |  | 408.373 | 0.783
|  | Serim BIOLIQ 2 | 45912 | 139 |  |  | 411.475 | 0.863
|  | Serim BIOLIQ 2 | 45862 | 41 |  |  | 401.244 | 0.784
PARACETAMOL 2, U/L | C6000-B:C2 | [9 empty] | 4.8694 | 4.626 | 5.1129 | 4.8694 | 4.626 | 5.1129
|  | RIAP L2-B | 2107EC | 280 |  |  | 319.991 | 4.752
|  | RIAP L2-B | 2228EC | 70 |  |  | 324.084 | 6.388
|  | RIAP L1-B | 2105EC | 288 |  |  | 152.374 | 4.181
|  | RIAP L2-B | 2124EC | 77 |  |  | 322.526 | 4.565
|  | RIAP L1-B | 2179EC | 123 |  |  | 154.079 | 5.751
PARATHYROID HORMONE, U/L | C6000-A:E2-1 | [9 empty] | 3.651 | 3.4684 | 3.8335 | 3.651 | 3.4684 | 3.8335
|  | PC V2 | 581988 | 55 |  |  | 20.859 | 1.373
|  | PC V1 | 612148 | 66 |  |  | 6.367 | 3.088
|  | PC V1 | 581987 | 55 |  |  | 6.467 | 4.138
|  | PC V2 | 573356 | 92 |  |  | 20.016 | 4.076
|  | PC V1 | 573355 | 92 |  |  | 6.138 | 4.493
|  | PC V1 | 595763 | 209 |  |  | 6.668 | 2.937
|  | PC V2 | 595765 | 235 |  |  | 19.985 | 4.048
PHOSPHATE, U/L | C6000-A:C1 | [9 empty] | 2.3143 | 2.1986 | 2.43 | 2.2882 | 2.1737 | 2.4026
|  | PCCC 2 | 461597 | 72 |  |  | 2.521 | 1.531
|  | PCCC 1 | 525027 | 417 |  |  | 1.29 | 2.974
|  | PCCC 2 | 519198 | 389 |  |  | 2.544 | 1.909
|  | PCCC 1 | 525205 | 164 |  |  | 1.159 | 2.05
|  | PCCC 2 | 535719 | 104 |  |  | 2.691 | 1.271
PHOSPHATE, U/L | C6000-B:C2 | [9 empty] | 2.2585 | 2.1455 | 2.3714 | 2.2882 | 2.1737 | 2.4026
|  | PCCC 2-B | 535719 | 89 |  |  | 2.721 | 1.968
|  | PCCC 1-B | 525027 | 366 |  |  | 1.288 | 1.988
|  | PCCC 2-B | 519198 | 346 |  |  | 2.564 | 2.281
|  | PCCC 1-B | 525205 | 151 |  |  | 1.19 | 3.172
|  | PCCC 2-B | 461597 | 70 |  |  | 2.464 | 1.202
POTASSIUM, U/L | C6000-A:ISE1 | [9 empty] | 1.8253 | 1.7341 | 1.9166 | 1.7767 | 1.6879 | 1.8655
|  | s-PCCC 2 | 535719 | 192 |  |  | 7.393 | 1.631
|  | PCCC 1 | 525027 | 106 |  |  | 3.65 | 1.654
|  | PCCC 2 | 461597 | 88 |  |  | 6.833 | 1.19
|  | s-PCCC 1 | 525205 | 330 |  |  | 3.586 | 2.236
|  | s-PCCC 2 | 519198 | 717 |  |  | 7.149 | 1.669
|  | s-PCCC 1 | 525027 | 597 |  |  | 3.645 | 1.915
POTASSIUM, U/L | C6000-B:ISE1 | [9 empty] | 1.7358 | 1.649 | 1.8226 | 1.7767 | 1.6879 | 1.8655
|  | s-PCCC 2-B | 535719 | 190 |  |  | 7.306 | 1.328
|  | PCCC 2-B | 461597 | 92 |  |  | 6.841 | 0.81
|  | s-PCCC 1-B | 525027 | 565 |  |  | 3.639 | 1.833
|  | PCCC 1-B | 525027 | 368 |  |  | 3.654 | 1.601
|  | PCCC 2-B | 519198 | 280 |  |  | 7.125 | 1.289
|  | s-PCCC 2-B | 519198 | 686 |  |  | 7.113 | 1.56
|  | s-PCCC 1-B | 525205 | 293 |  |  | 3.547 | 2.665
POTASSIUM URINE, U/L | C6000-B:ISE1 | [9 empty] | 3.7207 | 3.5347 | 3.9068 | 3.7207 | 3.5347 | 3.9068
|  | Ran-Liq 2 | 1209UC | 247 |  |  | 29.103 | 3.281
|  | Ran-Liq 3 | 1210UC | 249 |  |  | 71.532 | 5.462
|  | Bio-liq 2 | 88162 | 137 |  |  | 67.437 | 2.765
|  | Bio-liq1 | 88161 | 144 |  |  | 30.823 | 1.968
|  | Ran-Liq 2 | 1237UC | 137 |  |  | 29.08 | 2.725
|  | Ran-Liq 3 | 1238UC | 149 |  |  | 72.903 | 3.66
PROCALCITONIN (SENSITIVE), U/L | C6000-A:E1-2 | [9 empty] | 3.6233 | 3.4422 | 3.8045 | 3.6456 | 3.4633 | 3.8278
|  | PC PCT1 | 577640 | 33 |  |  | 0.462 | 3.161
|  | PC PCT2 | 577646 | 33 |  |  | 8.808 | 4.033
PROCALCITONIN (SENSITIVE), U/L | C6000-A:E2-1 | [9 empty] | 3.6476 | 3.4652 | 3.83 | 3.6456 | 3.4633 | 3.8278
|  | PC PCT2 | 671599 | 107 |  |  | 9.08 | 2.113
|  | PC PCT2 | 614523 | 61 |  |  | 9.569 | 2.66
|  | PC PCT1 | 671597 | 108 |  |  | 0.491 | 2.904
|  | PC PCT1 | 614519 | 61 |  |  | 0.496 | 2.425
|  | PC PCT2 | 559763 | 182 |  |  | 9.419 | 4.027
|  | PC PCT1 | 559760 | 181 |  |  | 0.509 | 4.799
PROGESTERONE III, U/L | C6000-A:E2-2 | [9 empty] | 6.5456 | 6.2184 | 6.8729 | 6.5456 | 6.2184 | 6.8729
|  | RIAP L1 | 2105EC | 287 |  |  | 3.378 | 6.885
|  | RIAP L3 | 2181EC | 135 |  |  | 98.292 | 4.296
|  | RIAP L3 | 2147 | 148 |  |  | 95.094 | 3.648
|  | RIAP L3 | 2110EC | 136 |  |  | 89.168 | 2.824
|  | RIAP L1 | 2179EC | 123 |  |  | 3.357 | 11.538
PROLACTIN, U/L | C6000-A:E1-2 | [9 empty] | 5.0234 | 4.7722 | 5.2746 | 5.0234 | 4.7722 | 5.2746
|  | RIAP L3 | 2110EC | 137 |  |  | 51.014 | 3.153
|  | RIAP L3 | 2181EC | 133 |  |  | 69.197 | 5.708
|  | RIAP L3 | 2147 | 153 |  |  | 44.917 | 5.06
|  | RIAP L2 | 2107EC | 279 |  |  | 28.545 | 4.706
|  | RIAP L2 | 2124EC | 78 |  |  | 26.833 | 8.272
|  | RIAP L2 | 2228EC | 64 |  |  | 32.167 | 1.674
PROSTATE-SPECIFIC AG, U/L | C6000-A:E2-2 | [9 empty] | 6.3552 | 6.0374 | 6.6729 | 6.3552 | 6.0374 | 6.6729
|  | RIAP L2 | 2124EC | 99 |  |  | 12.504 | 3.75
|  | RIAP L2 | 2107EC | 340 |  |  | 9.828 | 7.41
|  | RIAP L2 | 2228EC | 77 |  |  | 12.125 | 2.93
|  | RIAP L1 | 2105EC | 350 |  |  | 1.584 | 7.166
|  | RIAP L1 | 2179EC | 146 |  |  | 1.588 | 3.81
PROTEIN CSF, U/L | C6000-A:C1 | [9 empty] | 3.43 | 3.2585 | 3.6015 | 3.43 | 3.2585 | 3.6015
|  | PNPUC TPC3 | 673292 | 41 |  |  | 0.214 | 3.396
|  | PNPUC TPC3 | 602712 | 62 |  |  | 0.214 | 3.389
|  | PPPUC TPC3 | 602710 | 165 |  |  | 1.488 | 1.389
|  | PPPUC TPC3 | 531181 | 35 |  |  | 1.43 | 1.081
|  | PNPUC TPC3 | 531179 | 79 |  |  | 0.211 | 2.808
|  | PPPUC TPC3 | 646927 | 132 |  |  | 1.52 | 1.718
|  | PNPUC TPC3 | 634291 | 30 |  |  | 0.212 | 11.011
|  | PNPUC TPC3 | 649811 | 186 |  |  | 0.21 | 3.89
|  | PPPUC TPC3 | 637091 | 122 |  |  | 1.616 | 2.991
PROTEIN URINE, U/L | C6000-A:C1 | [9 empty] | 3.9337 | 3.7371 | 4.1304 | 3.9337 | 3.7371 | 4.1304
|  | PP PUC | 64692700 | 92 |  |  | 1.264 | 1.92
|  | PP PUC | 637091 | 122 |  |  | 1.342 | 2.016
|  | PP PUC | 602710 | 162 |  |  | 1.239 | 2.046
|  | PN PUC | 602712 | 120 |  |  | 0.178 | 3.139
|  | Ran-Liq 3 | 1238UC | 51 |  |  | 0.739 | 3.91
|  | Ran-Liq 2 | 1237UC | 50 |  |  | 0.135 | 4.105
|  | Ran-Liq 2 | 1209UC | 66 |  |  | 0.137 | 7.484
|  | PN PUC | 531179 | 81 |  |  | 0.175 | 3.148
|  | Ran-Liq 3 | 1210UC | 66 |  |  | 0.769 | 3.056
|  | PN PUC | 649811 | 203 |  |  | 0.173 | 5.783
|  | PP PUC | 531181 | 38 |  |  | 1.252 | 0.951
RHD3, U/L | C6000-B:C2 | [9 empty] | 2.6478 | 2.5154 | 2.7802 | 2.6478 | 2.5154 | 2.7802
|  | PPA1c H | 563451 | 515 |  |  | 10.612 | 3.079
|  | PPA1c H | 509113 | 51 |  |  | 10.726 | 2.216
|  | PNA1c H | 509111 | 51 |  |  | 5.725 | 1.995
|  | PPA1c H | 617429 | 222 |  |  | 9.444 | 2.733
|  | PNA1c H | 617428 | 224 |  |  | 5.684 | 2.241
|  | PNA1c H | 548925 | 344 |  |  | 5.687 | 2.539
|  | PNA1c H | 563450 | 165 |  |  | 5.755 | 2.025
RHEUMATOID FACTOR, U/L | C6000-B:C2 | [9 empty] | 3.9009 | 3.7059 | 4.096 | 3.9009 | 3.7059 | 4.096
|  | RFCO2 | 590051 | 116 |  |  | 52.38 | 3.15
|  | RFCO1 | 661133 | 131 |  |  | 23.404 | 4.107
|  | RFCO2 | 661134 | 134 |  |  | 53.296 | 3.727
|  | RFCO1 | 590050 | 119 |  |  | 22.129 | 4.996
|  | RFCO1 | 542412 | 33 |  |  | 21.355 | 2.306
|  | RFCO2 | 542413 | 33 |  |  | 50.661 | 2.651
Renin, U/L | LIAISON XL | [9 empty] | 5.6133 | 5.3327 | 5.894 | 5.6133 | 5.3327 | 5.894
|  | RENIN L2 | 7530320 | 42 |  |  | 111.185 | 5.659
|  | RENIN L1 | 7520320 | 49 |  |  | 24.972 | 5.574
SERUM FOLATE III, U/L | C6000-A:E2-2 | [9 empty] | 13.0086 | 12.3582 | 13.659 | 13.0086 | 12.3582 | 13.659
|  | RIAP L2 | 2228EC | 88 |  |  | 12.459 | 31.268
|  | RIAP L3 | 2147 | 198 |  |  | 23.148 | 6.955
|  | RIAP L3 | 2110EC | 159 |  |  | 23.348 | 5.985
|  | RIAP L3 | 2181EC | 174 |  |  | 22.875 | 14.869
|  | RIAP L2 | 2124EC | 98 |  |  | 12.508 | 14.032
|  | RIAP L2 | 2107EC | 352 |  |  | 12.134 | 7.969
SEX HORMONE BINDING GLOBULIN, U/L | C6000-A:E2-2 | [9 empty] | 4.2315 | 4.02 | 4.4431 | 4.2315 | 4.02 | 4.4431
|  | RIAP L1 | 2179EC | 115 |  |  | 9.938 | 3.383
|  | RIAP L1 | 2105EC | 297 |  |  | 11.022 | 5
|  | RIAP L2 | 2124EC | 82 |  |  | 38.312 | 4.717
|  | RIAP L2 | 2107EC | 268 |  |  | 37.77 | 3.788
|  | RIAP L2 | 2228EC | 66 |  |  | 35.521 | 2.476
SODIUM, U/L | C6000-A:ISE1 | [9 empty] | 1.54 | 1.463 | 1.6169 | 1.7338 | 1.6471 | 1.8205
|  | s-PCCC 2 | 535719 | 193 |  |  | 138.373 | 1.367
|  | s-PCCC 1 | 525027 | 588 |  |  | 114.427 | 1.614
|  | PCCC 2 | 461597 | 88 |  |  | 133.534 | 1.362
|  | s-PCCC 1 | 525205 | 342 |  |  | 114.234 | 1.755
|  | PCCC 1 | 525027 | 106 |  |  | 113.896 | 1.57
|  | s-PCCC 2 | 519198 | 714 |  |  | 137.38 | 1.423
SODIUM, U/L | C6000-B:ISE1 | [9 empty] | 1.8771 | 1.7833 | 1.971 | 1.7338 | 1.6471 | 1.8205
|  | s-PCCC 1-B | 525205 | 302 |  |  | 113.228 | 1.899
|  | s-PCCC 2-B | 535719 | 188 |  |  | 137.287 | 4.199
|  | PCCC 2-B | 519198 | 279 |  |  | 136.867 | 1.356
|  | PCCC 1-B | 525027 | 369 |  |  | 114.388 | 1.521
|  | PCCC 2-B | 461597 | 94 |  |  | 133.968 | 1.113
|  | s-PCCC 2-B | 519198 | 697 |  |  | 136.933 | 1.476
|  | s-PCCC 1-B | 525027 | 560 |  |  | 114.418 | 1.565
SODIUM URINE, U/L | C6000-B:ISE1 | [9 empty] | 7.6652 | 7.282 | 8.0485 | 7.6652 | 7.282 | 8.0485
|  | Ran-Liq 2 | 1209UC | 241 |  |  | 70.676 | 14.005
|  | Ran-Liq 3 | 1210UC | 252 |  |  | 165.329 | 7.702
|  | Bio-liq 2 | 88162 | 164 |  |  | 167.335 | 2.107
|  | Ran-Liq 3 | 1238UC | 155 |  |  | 172.542 | 1.691
|  | Ran-Liq 2 | 1237UC | 138 |  |  | 84.92 | 2.681
|  | Bio-liq1 | 88161 | 165 |  |  | 77.57 | 2.528
TESTOSTERONE, U/L | C6000-A:E2-1 | [9 empty] | 5.965 | 5.6668 | 6.2633 | 5.965 | 5.6668 | 6.2633
|  | RIAP L3 | 2181EC | 129 |  |  | 27.49 | 3.726
|  | RIAP L2 | 2228EC | 63 |  |  | 14.948 | 3.181
|  | RIAP L3 | 2147 | 145 |  |  | 27.647 | 5.899
|  | RIAP L2 | 2107EC | 273 |  |  | 12.882 | 7.282
|  | RIAP L2 | 2124EC | 78 |  |  | 14.379 | 3.188
|  | RIAP L3 | 2110EC | 134 |  |  | 28.937 | 6.944
THYROGLOBULIN v2, U/L | C6000-A:E1-1 | [9 empty] | 7.1672 | 6.8088 | 7.5255 | 7.1672 | 6.8088 | 7.5255
|  | RIAP L3 | 2147 | 153 |  |  | 65.571 | 4.306
|  | RIAP L1 | 2105EC | 302 |  |  | 2.099 | 5.159
|  | RIAP L3 | 2181EC | 130 |  |  | 62.073 | 10.34
|  | RIAP L3 | 2110EC | 144 |  |  | 67.02 | 4.727
|  | RIAP L1 | 2179EC | 134 |  |  | 1.939 | 11.056
THYROID STIMULATING HORMONE, U/L | C6000-A:E2-2 | [9 empty] | 5.3614 | 5.0934 | 5.6295 | 5.5211 | 5.245 | 5.7972
|  | RIAP L2 | 2228EC | 98 |  |  | 2.595 | 2.576
|  | RIAP L2 | 2107EC | 420 |  |  | 2.399 | 6.018
|  | RIAP L2 | 2124EC | 119 |  |  | 2.93 | 4.83
|  | RIAP L1 | 2179EC | 174 |  |  | 0.133 | 6.438
|  | RIAP L1 | 2105EC | 436 |  |  | 0.103 | 4.785
THYROID STIMULATING HORMONE, U/L | C6000-B:E3-1 | [9 empty] | 5.7272 | 5.4408 | 6.0135 | 5.5211 | 5.245 | 5.7972
|  | RIAP L2-B | 2124EC | 90 |  |  | 2.965 | 7.416
|  | RIAP L1-B | 2179EC | 141 |  |  | 0.136 | 6.233
|  | RIAP L1-B | 2105EC | 316 |  |  | 0.106 | 5.018
|  | RIAP L2-B | 2107EC | 312 |  |  | 2.392 | 5.914
|  | RIAP L2-B | 2228EC | 77 |  |  | 2.591 | 4.283
TOTAL BILIRUBIN, U/L | C6000-A:C1 | [9 empty] | 4.1958 | 3.986 | 4.4056 | 4.2068 | 3.9964 | 4.4171
|  | MAS BILIRUBIN 3 | BC23063A | 73 |  |  | 321.348 | 3.499
|  | PCCC 2 | 519198 | 417 |  |  | 60.758 | 3.822
|  | PCCC 2 | 535719 | 108 |  |  | 63.277 | 3.941
|  | PCCC 2 | 461597 | 77 |  |  | 61.884 | 3.082
|  | PCCC 1 | 525205 | 168 |  |  | 16.461 | 4.504
|  | PCCC 1 | 525027 | 451 |  |  | 15.771 | 5.593
|  | MAS BILIRUBIN 3 | BC24033A | 414 |  |  | 346.205 | 2.757
TOTAL BILIRUBIN, U/L | C6000-B:C2 | [9 empty] | 4.2184 | 4.0074 | 4.4293 | 4.2068 | 3.9964 | 4.4171
|  | PCCC 2-B | 519198 | 397 |  |  | 60.03 | 4.845
|  | PCCC 2-B | 535719 | 99 |  |  | 61.335 | 2.977
|  | PCCC 1-B | 525205 | 166 |  |  | 16.427 | 4.649
|  | PCCC 2-B | 461597 | 78 |  |  | 62.617 | 1.79
|  | MAS BILIRUBIN-B | BC23063A | 72 |  |  | 324.158 | 2.844
|  | MAS BILIRUBIN-B | BC24033A | 385 |  |  | 342.326 | 3.409
|  | PCCC 1-B | 525027 | 418 |  |  | 15.66 | 4.798
TOTAL CHOLESTEROL, U/L | C6000-A:C1 | [9 empty] | 2.2261 | 2.1148 | 2.3374 | 2.1084 | 2.0029 | 2.2138
|  | PCCC 1 | 525027 | 93 |  |  | 2.378 | 2.353
|  | s-PCCC 2 | 519198 | 610 |  |  | 4.467 | 1.944
|  | s-PCCC 1 | 525205 | 264 |  |  | 2.427 | 2.244
|  | s-PCCC 2 | 535719 | 165 |  |  | 4.347 | 2.356
|  | s-PCCC 1 | 525027 | 521 |  |  | 2.424 | 2.454
|  | PCCC 2 | 461597 | 76 |  |  | 4.485 | 2.195
TOTAL CHOLESTEROL, U/L | C6000-B:C2 | [9 empty] | 1.9771 | 1.8783 | 2.076 | 2.1084 | 2.0029 | 2.2138
|  | s-PCCC 1-B | 525205 | 239 |  |  | 2.423 | 2.401
|  | s-PCCC 2-B | 519198 | 604 |  |  | 4.49 | 1.966
|  | PCCC 2-B | 461597 | 64 |  |  | 4.538 | 1.289
|  | PCCC 1-B | 525027 | 76 |  |  | 2.399 | 1.686
|  | s-PCCC 2-B | 535719 | 158 |  |  | 4.31 | 1.663
|  | s-PCCC 1-B | 525027 | 505 |  |  | 2.432 | 1.97
TOTAL PROTEIN, U/L | C6000-A:C1 | [9 empty] | 1.9798 | 1.8808 | 2.0788 | 1.997 | 1.8972 | 2.0969
|  | PCCC 1 | 525027 | 408 |  |  | 48.055 | 1.616
|  | PCCC 2 | 461597 | 73 |  |  | 80.881 | 1.34
|  | PCCC 2 | 519198 | 387 |  |  | 78.18 | 1.584
|  | PCCC 2 | 535719 | 99 |  |  | 75.53 | 3.091
|  | PCCC 1 | 525205 | 164 |  |  | 48.256 | 2.842
TOTAL PROTEIN, U/L | C6000-B:C2 | [9 empty] | 2.018 | 1.9171 | 2.1189 | 1.997 | 1.8972 | 2.0969
|  | s-PCCC 1-B | 525027 | 42 |  |  | 47.695 | 1.12
|  | PCCC 1-B | 525205 | 151 |  |  | 47.27 | 2.31
|  | PCCC 2-B | 519198 | 327 |  |  | 77.293 | 2.019
|  | s-PCCC 2-B | 461597 | 38 |  |  | 80.203 | 1.087
|  | PCCC 2-B | 535719 | 89 |  |  | 74.152 | 2.365
|  | PCCC 1-B | 525027 | 277 |  |  | 47.584 | 1.924
TRANSFERRIN, U/L | C6000-A:C1 | [9 empty] | 2.9476 | 2.8002 | 3.0949 | 2.9476 | 2.8002 | 3.0949
|  | s-PCCC 1 | 525027 | 54 |  |  | 1.954 | 3.332
|  | s-PCCC 2 | 461597 | 48 |  |  | 3.392 | 3.989
|  | PCCC 2 | 535719 | 110 |  |  | 3.356 | 1.794
|  | PCCC 2 | 519198 | 425 |  |  | 3.446 | 3.288
|  | PCCC 1 | 525205 | 188 |  |  | 1.919 | 1.867
|  | PCCC 1 | 525027 | 365 |  |  | 2.014 | 3.04
TRI-IODO THYRONINE, U/L | C6000-A:E1-2 | [9 empty] | 4.4766 | 4.2528 | 4.7004 | 4.4766 | 4.2528 | 4.7004
|  | RIAP L2 | 2228EC | 65 |  |  | 15.641 | 2.536
|  | RIAP L2 | 2107EC | 276 |  |  | 13.81 | 3.468
|  | RIAP L2 | 2124EC | 74 |  |  | 15.164 | 3.567
|  | RIAP L1 | 2179EC | 117 |  |  | 5.052 | 4.878
|  | RIAP L1 | 2105EC | 294 |  |  | 4.403 | 5.556
TRIGLYCERIDE, U/L | C6000-A:C1 | [9 empty] | 2.5894 | 2.4599 | 2.7189 | 2.5894 | 2.4599 | 2.7189
|  | PCCC 1 | 525027 | 341 |  |  | 1.35 | 3.507
|  | PCCC 2 | 519198 | 385 |  |  | 2.453 | 1.935
|  | PCCC 2 | 535719 | 107 |  |  | 2.444 | 2.139
|  | s-PCCC 1 | 525027 | 52 |  |  | 1.341 | 2.249
|  | s-PCCC 2 | 461597 | 42 |  |  | 2.443 | 1.869
|  | PCCC 1 | 525205 | 166 |  |  | 1.343 | 2.178
UREA, U/L | C6000-A:C1 | [9 empty] | 2.7265 | 2.5902 | 2.8629 | 2.6416 | 2.5095 | 2.7736
|  | s-PCCC 1 | 525027 | 546 |  |  | 6.708 | 2.652
|  | PCCC 2 | 461597 | 76 |  |  | 19.705 | 2.134
|  | s-PCCC 2 | 519198 | 644 |  |  | 20.406 | 2.645
|  | s-PCCC 2 | 535719 | 167 |  |  | 19.949 | 3.139
|  | s-PCCC 1 | 525205 | 267 |  |  | 6.6 | 3.154
|  | PCCC 1 | 525027 | 95 |  |  | 6.636 | 1.915
UREA, U/L | C6000-B:C2 | [9 empty] | 2.5518 | 2.4242 | 2.6794 | 2.6416 | 2.5095 | 2.7736
|  | PCCC 2-B | 461597 | 77 |  |  | 19.677 | 1.878
|  | PCCC 1-B | 525027 | 91 |  |  | 6.56 | 1.953
|  | s-PCCC 1-B | 525027 | 530 |  |  | 6.778 | 2.666
|  | s-PCCC 1-B | 525205 | 256 |  |  | 6.739 | 2.981
|  | s-PCCC 2-B | 519198 | 639 |  |  | 20.671 | 2.502
|  | s-PCCC 2-B | 535719 | 164 |  |  | 20.051 | 2.175
UREA URINE, U/L | C6000-A:C1 | [9 empty] | 10.3063 | 9.791 | 10.8216 | 5.8892 | 5.5947 | 6.1837
|  | Ran-Liq 2 | 1237UC | 50 |  |  | 149.36 | 3.865
|  | Ran-Liq 3 | 1238UC | 49 |  |  | 281.529 | 4.269
|  | Ran-Liq 3 | 1210UC | 73 |  |  | 275.344 | 7.779
|  | Ran-Liq 2 | 1209UC | 73 |  |  | 153.532 | 16.514
UREA URINE, U/L | C6000-B:C2 | [9 empty] | 4.0497 | 3.8472 | 4.2522 | 5.8892 | 5.5947 | 6.1837
|  | Ran-Liq 2 | 1209UC | 194 |  |  | 149.709 | 3.381
|  | Bio-liq1 | 88161 | 136 |  |  | 151.508 | 3.45
|  | Ran-Liq 3 | 1210UC | 229 |  |  | 278.159 | 5.411
|  | Bio-liq 2 | 88162 | 131 |  |  | 257.644 | 3.843
|  | Ran-Liq 2 | 1237UC | 122 |  |  | 149.59 | 3.724
|  | Ran-Liq 3 | 1238UC | 137 |  |  | 285.408 | 3.184
URIC ACID, U/L | C6000-A:C1 | [9 empty] | 2.1078 | 2.0024 | 2.2131 | 2.046 | 1.9437 | 2.1483
|  | PCCC 1 | 525205 | 141 |  |  | 0.287 | 2.129
|  | PCCC 1 | 525027 | 357 |  |  | 0.289 | 2.198
|  | PCCC 2 | 461597 | 67 |  |  | 0.571 | 1.574
|  | PCCC 2 | 519198 | 326 |  |  | 0.588 | 2.119
|  | PCCC 2 | 535719 | 93 |  |  | 0.605 | 2.011
URIC ACID, U/L | C6000-B:C2 | [9 empty] | 1.971 | 1.8724 | 2.0695 | 2.046 | 1.9437 | 2.1483
|  | PCCC 2-B | 535719 | 79 |  |  | 0.596 | 2.095
|  | PCCC 1-B | 525027 | 255 |  |  | 0.287 | 2.063
|  | s-PCCC 2-B | 461597 | 33 |  |  | 0.569 | 0.922
|  | PCCC 1-B | 525205 | 134 |  |  | 0.285 | 2.003
|  | PCCC 2-B | 519198 | 300 |  |  | 0.584 | 2.009
|  | s-PCCC 1-B | 525027 | 38 |  |  | 0.286 | 1.104
URIC ACID URINE, U/L | C6000-B:C2 | [9 empty] | 4.5006 | 4.2756 | 4.7257 | 4.5006 | 4.2756 | 4.7257
|  | Bio-liq1 | 88161 | 104 |  |  | 0.854 | 2.474
|  | Bio-liq 2 | 88162 | 99 |  |  | 1.282 | 2.279
|  | Ran-Liq 3 | 1238UC | 117 |  |  | 0.871 | 2.163
|  | Ran-Liq 2 | 1237UC | 108 |  |  | 0.545 | 2.637
|  | Ran-Liq 3 | 1210UC | 198 |  |  | 0.861 | 4.613
|  | Ran-Liq 2 | 1209UC | 197 |  |  | 0.541 | 7.118
Urine osmolality, U/L | OSMO1 | [9 empty] | 0.7054 | 0.67013 | 0.74067 | 0.7054 | 0.67013 | 0.74067
|  | Biorad Urine QC1 | 88162 | 93 |  |  | 787.753 | 0.411
|  | Ran-Liq 2 | 1237UC | 79 |  |  | 436.139 | 1.175
|  | Biorad Urine QC1 | 88161 | 102 |  |  | 436.5 | 0.438
|  | Randox Urine QC 3 | 1210UC | 34 |  |  | 807.118 | 0.585
|  | Ran-Liq 3 | 1238UC | 87 |  |  | 828.839 | 0.694
|  | Randox Urine QC 2 | 1209UC | 32 |  |  | 415.156 | 0.676
VITAMIN B12 II, U/L | C6000-A:E1-2 | [9 empty] | 6.188 | 5.8786 | 6.4974 | 6.188 | 5.8786 | 6.4974
|  | RIAP L2 | 2228EC | 91 |  |  | 443.052 | 3.455
|  | RIAP L1 | 2105EC | 436 |  |  | 171.022 | 7.437
|  | RIAP L2 | 2107EC | 407 |  |  | 441.334 | 4.036
|  | RIAP L1 | 2179EC | 190 |  |  | 159.453 | 6.255
|  | RIAP L2 | 2124EC | 115 |  |  | 436.328 | 8.434
proBNP, U/L | C6000-A:E1-1 | [9 empty] | 3.5861 | 3.4068 | 3.7654 | 3.5861 | 3.4068 | 3.7654
|  | PC CARD1 | 657390 | 65 |  |  | 126.177 | 4.313
|  | PC CARD2 | 657389 | 63 |  |  | 4484.127 | 4.1
|  | PC CARD1 | 629283 | 63 |  |  | 128.321 | 3.826
|  | PC CARD1 | 586866 | 87 |  |  | 121.037 | 3.93
|  | PC CARD2 | 629284 | 61 |  |  | 4743.607 | 3.303
|  | PC CARD1 | 623415 | 163 |  |  | 123.507 | 3.259
|  | PC CARD2 | 623416 | 162 |  |  | 4727.735 | 2.991
|  | PC CARD2 | 586867 | 87 |  |  | 4871.494 | 3.844

## Raw Data
Instrument | Test | QC | QC lot No. | Target mean | Target SD | SDI | Number of values | Bias | Calculated mean | Calculated SD | CV% | Usage
C6000-B:C2 | ADENOSINE DEAMINASE FLUID | ADA 2 | 201220 | 141 | 8 | 0.193 | 132 | 0.011 | 142.541 | 6.342 | 4.45 | In use
C6000-B:C2 | ADENOSINE DEAMINASE FLUID | ADA 2 | ADC203210 | 144 | 8.2 | -0.445 | 46 | -0.025 | 140.348 | 6.5 | 4.632 | Study
C6000-B:C2 | ADENOSINE DEAMINASE FLUID | ADA 1 | ADC103210 | 30.8 | 2.6 | 0.365 | 117 | 0.031 | 31.75 | 1.556 | 4.901 | Study
C6000-B:C2 | ADENOSINE DEAMINASE FLUID | ADA 2 | 202210 | 151.4 | 8.56 | 0.484 | 102 | 0.027 | 155.545 | 5.983 | 3.846 | Study
C6000-B:C2 | ADENOSINE DEAMINASE FLUID | ADA 0 | 01220 | 11.3 | 0.6 | 0.263 | 131 | 0.014 | 11.458 | 0.612 | 5.337 | In use
C6000-B:C2 | ADENOSINE DEAMINASE FLUID | ADA 0 | ADCL02210 | 10.7 | 0.6 | 0.932 | 106 | 0.052 | 11.259 | 0.343 | 3.044 | Study
C6000-B:C2 | ADENOSINE DEAMINASE FLUID | ADA 1 | 101220 | 29.8 | 1.8 | 0.365 | 96 | 0.022 | 30.457 | 1.223 | 4.017 | In use
C6000-B:C2 | ADENOSINE DEAMINASE FLUID | ADA 1 | 102210 | 31.1 | 1.76 | 0.818 | 67 | 0.046 | 32.54 | 0.964 | 2.964 | Study
C6000-B:C2 | ADENOSINE DEAMINASE FLUID | ADA 0 | L01210 | 10.2 | 0.567 | 0.878 | 49 | 0.049 | 10.698 | 0.297 | 2.775 | Study
C6000-A:E1-1 | ADRENOCORTICOTROPIC HORMONE | PC MM1 | 523425 | 8.87 | 0.89 | -0.055 | 40 | -0.006 | 8.821 | 0.456 | 5.165 | Study
C6000-A:E1-1 | ADRENOCORTICOTROPIC HORMONE | PC MM2 | 523426 | 205 | 10.3 | -1.031 | 39 | -0.052 | 194.382 | 9.222 | 4.744 | Study
C6000-A:C1 | ALANINE TRANSAMINASE | PCCC 1 | 525027 | 47.2 | 2.8 | -0.468 | 329 | -0.028 | 45.889 | 1.346 | 2.932 | Study
C6000-A:C1 | ALANINE TRANSAMINASE | PCCC 2 | 535719 | 127 | 8 | -0.614 | 93 | -0.039 | 122.088 | 3.398 | 2.783 | In use
C6000-B:C2 | ALANINE TRANSAMINASE | PCCC 1-B | 525205 | 48.9 | 2.9 | -0.323 | 158 | -0.019 | 47.963 | 1.331 | 2.775 | In use
C6000-A:C1 | ALANINE TRANSAMINASE | s-PCCC 2 | 461597 | 122 | 7 | -0.008 | 39 | 0 | 121.946 | 2.565 | 2.104 | Study
C6000-B:C2 | ALANINE TRANSAMINASE | s-PCCC 2-B | 461597 | 122 | 7 | -0.476 | 41 | -0.027 | 118.668 | 1.439 | 1.212 | Study
C6000-A:C1 | ALANINE TRANSAMINASE | PCCC 2 | 519198 | 120 | 7 | -0.474 | 369 | -0.028 | 116.679 | 2.825 | 2.422 | Study
C6000-B:C2 | ALANINE TRANSAMINASE | PCCC 2-B | 535719 | 127 | 8 | -0.587 | 98 | -0.037 | 122.306 | 12.912 | 10.557 | In use
C6000-B:C2 | ALANINE TRANSAMINASE | PCCC 2-B | 519198 | 120 | 7 | -0.548 | 365 | -0.032 | 116.165 | 4.581 | 3.943 | Study
C6000-A:C1 | ALANINE TRANSAMINASE | s-PCCC 1 | 525027 | 47.2 | 2.8 | 0.091 | 47 | 0.005 | 47.455 | 1.308 | 2.757 | Study
C6000-B:C2 | ALANINE TRANSAMINASE | s-PCCC 1-B | 525027 | 47.2 | 2.8 | -0.393 | 46 | -0.023 | 46.1 | 0.705 | 1.53 | Study
C6000-A:C1 | ALANINE TRANSAMINASE | PCCC 1 | 525205 | 48.9 | 2.9 | -0.444 | 147 | -0.026 | 47.612 | 0.918 | 1.929 | In use
C6000-B:C2 | ALANINE TRANSAMINASE | PCCC 1-B | 525027 | 47.2 | 2.8 | -0.657 | 311 | -0.039 | 45.36 | 1.821 | 4.014 | Study
C6000-B:C2 | ALBUMIN | s-PCCC 2-B | 519198 | 50.2 | 3 | -0.083 | 584 | -0.005 | 49.95 | 1.388 | 2.779 | Study
C6000-A:C1 | ALBUMIN | PCCC 1 | 525205 | 32.1 | 1.9 | 0.223 | 141 | 0.013 | 32.523 | 1.04 | 3.197 | In use
C6000-B:C2 | ALBUMIN | s-PCCC 1-B | 525205 | 32.1 | 1.9 | -0.093 | 227 | -0.005 | 31.924 | 1.378 | 4.315 | In use
C6000-B:C2 | ALBUMIN | s-PCCC 2-B | 535719 | 48.8 | 2.9 | -0.403 | 146 | -0.024 | 47.631 | 1.104 | 2.319 | In use
C6000-B:C2 | ALBUMIN | PCCC 2-B | 461597 | 51.2 | 1.6 | 0.392 | 68 | 0.012 | 51.828 | 0.965 | 1.861 | Study
C6000-B:C2 | ALBUMIN | s-PCCC 1-B | 525027 | 31.9 | 1.6 | -0.067 | 488 | -0.003 | 31.792 | 1.012 | 3.183 | Study
C6000-A:C1 | ALBUMIN | PCCC 2 | 519198 | 50.2 | 3 | -0.137 | 371 | -0.008 | 49.789 | 0.912 | 1.831 | Study
C6000-A:C1 | ALBUMIN | PCCC 2 | 535719 | 48.8 | 2.9 | -0.069 | 88 | -0.004 | 48.601 | 1.052 | 2.165 | In use
C6000-A:C1 | ALBUMIN | PCCC 2 | 461597 | 51.2 | 1.6 | 0.497 | 73 | 0.016 | 51.996 | 1.047 | 2.014 | Study
C6000-A:C1 | ALBUMIN | PCCC 1 | 525027 | 31.9 | 1.6 | -0.186 | 407 | -0.009 | 31.602 | 0.912 | 2.886 | Study
C6000-B:C2 | ALBUMIN | PCCC 1-B | 525027 | 31.9 | 1.6 | -0.584 | 78 | -0.029 | 30.965 | 0.66 | 2.133 | Study
C6000-A:C1 | ALBUMIN CSF | LIQ SFC2 | 56502 | 348 | 16.8 | -0.335 | 110 | -0.016 | 342.374 | 11.761 | 3.435 | Study
C6000-A:C1 | ALBUMIN CSF | LIQ SFC2 | 56492 | 346 | 16.7 | 0.279 | 48 | 0.013 | 350.658 | 6.191 | 1.765 | Study
C6000-A:C1 | ALBUMIN CSF | PPPUC TPC3 | 646927 | 116 | 9 | 0.573 | 121 | 0.044 | 121.158 | 4.841 | 3.996 | In use
C6000-A:C1 | ALBUMIN CSF | LIQ SFC2 | 56522 | 323 | 22.5 | -0.224 | 128 | -0.016 | 317.951 | 11.625 | 3.656 | Study
C6000-A:C1 | ALBUMIN CSF | PPPUC TPC3 | 637091 | 120 | 10 | 0.178 | 120 | 0.015 | 121.775 | 6.003 | 4.929 | Study
C6000-A:C1 | ALBUMIN CSF | PPPUC TPC3 | 531181 | 122 | 10 | -0.233 | 37 | -0.019 | 119.67 | 3.483 | 2.911 | Study
C6000-A:C1 | ALBUMIN CSF | LIQ SFC2 | 56532 | 333 | 23.25 | -1.144 | 39 | -0.08 | 306.392 | 9.012 | 2.941 | In use
C6000-A:C1 | ALBUMIN CSF | PPPUC TPC3 | 602710 | 129 | 10 | -0.125 | 157 | -0.01 | 127.754 | 5.535 | 4.332 | Study
C6000-A:C1 | ALBUMIN CSF | LIQ SFC2 | 56512 | 359 | 17.3 | -0.317 | 100 | -0.015 | 353.524 | 10.198 | 2.885 | Study
C6000-A:C1 | ALBUMIN URINE | PN PUC | 602712 | 30.8 | 2.5 | 0.442 | 123 | 0.036 | 31.905 | 1.204 | 3.774 | Study
C6000-A:C1 | ALBUMIN URINE | PP PUC | 64692700 | 116 | 9 | -0.208 | 84 | -0.016 | 114.132 | 2.784 | 2.439 | Study
C6000-A:C1 | ALBUMIN URINE | PN PUC | 649811 | 30.2 | 2.4 | 0.694 | 183 | 0.055 | 31.864 | 6.019 | 18.891 | Study
C6000-A:C1 | ALBUMIN URINE | Ran-Liq 2 | 1237UC | 24 | 2.4 | 0.463 | 75 | 0.046 | 25.111 | 5.751 | 22.904 | In use
C6000-A:C1 | ALBUMIN URINE | PN PUC | 531179 | 27.7 | 2.2 | 0.108 | 82 | 0.009 | 27.938 | 1.007 | 3.603 | Study
C6000-A:C1 | ALBUMIN URINE | PP PUC | 531181 | 116 | 9 | -0.191 | 40 | -0.015 | 114.28 | 2.08 | 1.82 | Study
C6000-A:C1 | ALBUMIN URINE | Ran-Liq 2 | 1209UC | 24.1 | 2.4 | -0.163 | 68 | -0.016 | 23.709 | 2.181 | 9.198 | Study
C6000-A:C1 | ALBUMIN URINE | Ran-Liq 3 | 1238UC | 62.4 | 6.25 | 0.12 | 79 | 0.012 | 63.152 | 4.03 | 6.381 | In use
C6000-A:C1 | ALBUMIN URINE | Ran-Liq 3 | 1210UC | 64.1 | 6.4 | -0.096 | 68 | -0.01 | 63.488 | 2.041 | 3.215 | Study
C6000-A:C1 | ALBUMIN URINE | PP PUC | 637091 | 117 | 9 | 0.33 | 121 | 0.025 | 119.967 | 4.208 | 3.508 | Study
C6000-A:C1 | ALBUMIN URINE | PP PUC | 602710 | 121 | 10 | 0.105 | 161 | 0.009 | 122.05 | 4.113 | 3.37 | Study
LIAISON XL | ALDO | ALDO L1 | 227012 | 181.44 | 18.19 | 0.432 | 51 | 0.043 | 189.294 | 18.937 | 10.004 | In use
LIAISON XL | ALDO | ALDO L2 | 228012 | 720 | 73 | 0.279 | 48 | 0.028 | 740.333 | 45.766 | 6.182 | In use
C6000-B:C2 | ALKALINE PHOSPHATASE | s-PCCC 2-B | 519198 | 236 | 15 | 0.118 | 588 | 0.008 | 237.776 | 10.384 | 4.367 | Study
C6000-A:C1 | ALKALINE PHOSPHATASE | s-PCCC 2 | 519198 | 236 | 15 | 0.254 | 619 | 0.016 | 239.817 | 10.147 | 4.231 | Study
C6000-B:C2 | ALKALINE PHOSPHATASE | PCCC 2-B | 461597 | 231 | 14 | -0.055 | 68 | -0.003 | 230.235 | 6.386 | 2.774 | Study
C6000-B:C2 | ALKALINE PHOSPHATASE | s-PCCC 1-B | 525205 | 101 | 6 | -0.598 | 254 | -0.036 | 97.413 | 6.403 | 6.573 | In use
C6000-A:C1 | ALKALINE PHOSPHATASE | PCCC 2 | 461597 | 231 | 14 | 0.045 | 69 | 0.003 | 231.623 | 6.793 | 2.933 | Study
C6000-A:C1 | ALKALINE PHOSPHATASE | s-PCCC 1 | 525205 | 101 | 6 | -0.035 | 268 | -0.002 | 100.791 | 5.567 | 5.523 | In use
C6000-B:C2 | ALKALINE PHOSPHATASE | s-PCCC 1-B | 525027 | 91.5 | 7 | -0.153 | 484 | -0.012 | 90.426 | 4.361 | 4.822 | Study
C6000-A:C1 | ALKALINE PHOSPHATASE | PCCC 1 | 525027 | 91.5 | 7 | -0.155 | 92 | -0.012 | 90.413 | 4.201 | 4.647 | Study
C6000-B:C2 | ALKALINE PHOSPHATASE | PCCC 1-B | 525027 | 91.5 | 7 | -0.475 | 80 | -0.036 | 88.175 | 4.505 | 5.11 | Study
C6000-A:C1 | ALKALINE PHOSPHATASE | s-PCCC 1 | 525027 | 91.5 | 7 | -0.06 | 520 | -0.005 | 91.083 | 4.835 | 5.308 | Study
C6000-A:C1 | ALKALINE PHOSPHATASE | s-PCCC 2 | 535719 | 267 | 16 | 0.013 | 160 | 0.001 | 267.206 | 14.618 | 5.471 | In use
C6000-B:C2 | ALKALINE PHOSPHATASE | s-PCCC 2-B | 535719 | 267 | 16 | -0.558 | 162 | -0.033 | 258.068 | 10.761 | 4.17 | In use
C6000-B:C2 | ALPHA 1-ANTITRYPSIN | PCCC 1-B | 525027 | 0.802 | 0.06 | -0.004 | 231 | 0 | 0.802 | 0.024 | 3.023 | Study
C6000-B:C2 | ALPHA 1-ANTITRYPSIN | PCCC 1-B | 525205 | 0.883 | 0.071 | 0.301 | 139 | 0.024 | 0.904 | 0.022 | 2.436 | In use
C6000-B:C2 | ALPHA 1-ANTITRYPSIN | s-PCCC 2-B | 461597 | 1.23 | 0.1 | 0.474 | 33 | 0.039 | 1.277 | 0.02 | 1.568 | Study
C6000-B:C2 | ALPHA 1-ANTITRYPSIN | s-PCCC 1-B | 525027 | 0.802 | 0.06 | 0.004 | 36 | 0 | 0.802 | 0.014 | 1.738 | Study
C6000-B:C2 | ALPHA 1-ANTITRYPSIN | PCCC 2-B | 519198 | 1.32 | 0.06 | -0.525 | 277 | -0.024 | 1.289 | 0.037 | 2.863 | Study
C6000-B:C2 | ALPHA 1-ANTITRYPSIN | PCCC 2-B | 535719 | 1.32 | 0.11 | 0.209 | 85 | 0.017 | 1.343 | 0.019 | 1.409 | In use
C6000-A:E1-1 | ALPHA-FETO PROTEIN | RIAP L3 | 2147 | 191 | 19 | 2.129 | 158 | 0.212 | 231.45 | 15.3 | 6.611 | Study
C6000-A:E1-1 | ALPHA-FETO PROTEIN | RIAP L3 | 2181EC | 218 | 22 | -0.555 | 145 | -0.056 | 205.785 | 8.435 | 4.099 | In use
C6000-A:E1-1 | ALPHA-FETO PROTEIN | RIAP L1 | 2179EC | 10.6 | 1.04 | -0.398 | 139 | -0.039 | 10.187 | 0.605 | 5.934 | In use
C6000-A:E1-1 | ALPHA-FETO PROTEIN | RIAP L1 | 2105EC | 11.3 | 1.24 | 0.335 | 311 | 0.037 | 11.715 | 0.56 | 4.779 | Study
C6000-A:E1-1 | ALPHA-FETO PROTEIN | RIAP L3 | 2110EC | 226 | 22 | 0.254 | 149 | 0.025 | 231.58 | 8.7 | 3.757 | Study
C6000-B:C2 | AMMONIA | AMM-P | 623128 | 231 | 15 | 0.183 | 81 | 0.012 | 233.749 | 4.935 | 2.111 | Study
C6000-B:C2 | AMMONIA | AMM-N | 644004 | 62.5 | 4.2 | 0.243 | 167 | 0.016 | 63.522 | 2.299 | 3.62 | In use
C6000-B:C2 | AMMONIA | AMM-N | 587880 | 61.6 | 4.1 | 0.464 | 232 | 0.031 | 63.501 | 2.631 | 4.143 | Study
C6000-B:C2 | AMMONIA | AMM-P | 587881 | 219 | 15 | 0.315 | 227 | 0.022 | 223.73 | 6.396 | 2.859 | Study
C6000-B:C2 | AMMONIA | AMM-P | 644006 | 228 | 15 | -0.01 | 87 | -0.001 | 227.843 | 6.438 | 2.825 | In use
C6000-A:E2-1 | ANTI-THYROGLOBULIN AB | PC THYRO1 | 580444 | 113 | 11.3 | -0.029 | 228 | -0.003 | 112.67 | 8.097 | 7.186 | Study
C6000-A:E2-1 | ANTI-THYROGLOBULIN AB | PC THYRO2 | 580446 | 224 | 22.4 | -0.103 | 231 | -0.01 | 221.702 | 16.968 | 7.653 | Study
C6000-A:E2-1 | ANTI-THYROGLOBULIN AB | PC THYRO2 | 612140 | 221 | 22.1 | -0.147 | 110 | -0.015 | 217.754 | 12.418 | 5.703 | Study
C6000-A:E2-1 | ANTI-THYROGLOBULIN AB | PC THYRO1 | 612139 | 88.8 | 8.88 | 1.405 | 87 | 0.14 | 101.275 | 3.948 | 3.899 | Study
C6000-A:E2-1 | ANTI-THYROGLOBULIN AB | PC THYRO1 | 669238 | 58 | 5.8 | 0.791 | 49 | 0.079 | 62.587 | 2.929 | 4.679 | In use
C6000-A:E1-1 | ANTI-THYROID PEROXIDASE AB | PC THYRO2 | 580446 | 109 | 7.63 | 0.054 | 42 | 0.004 | 109.415 | 8.373 | 7.653 | Study
C6000-A:E1-1 | ANTI-THYROID PEROXIDASE AB | PC THYRO1 | 580444 | 46.1 | 4.61 | 0.353 | 43 | 0.035 | 47.726 | 4.714 | 9.876 | Study
C6000-B:C2 | APOPROTEIN A | PCCC 1-B | 525027 | 1.1 | 0.09 | 0.212 | 98 | 0.017 | 1.119 | 0.049 | 4.366 | Study
C6000-B:C2 | APOPROTEIN A | PCCC 2-B | 519198 | 1.67 | 0.13 | 0.592 | 116 | 0.046 | 1.747 | 0.079 | 4.51 | Study
C6000-B:C2 | APOPROTEIN A | PCCC 1-B | 525205 | 1.1 | 0.09 | 0.59 | 38 | 0.048 | 1.153 | 0.057 | 4.952 | In use
C6000-B:C2 | APOPROTEIN B | PCCC 1-B | 525205 | 0.484 | 0.039 | -0.239 | 39 | -0.019 | 0.475 | 0.026 | 5.424 | In use
C6000-B:C2 | APOPROTEIN B | PCCC 2-B | 519198 | 0.779 | 0.062 | -0.239 | 110 | -0.019 | 0.764 | 0.037 | 4.809 | Study
C6000-B:C2 | APOPROTEIN B | PCCC 1-B | 525027 | 0.481 | 0.03 | -0.345 | 93 | -0.022 | 0.471 | 0.024 | 5.087 | Study
C6000-A:C1 | ASPARTATE TRANSAMINASE | PCCC 1 | 525205 | 44.7 | 2.7 | -0.154 | 141 | -0.009 | 44.284 | 0.727 | 1.641 | In use
C6000-A:C1 | ASPARTATE TRANSAMINASE | PCCC 2 | 535719 | 143 | 9 | -0.605 | 91 | -0.038 | 137.558 | 2.425 | 1.763 | In use
C6000-A:C1 | ASPARTATE TRANSAMINASE | s-PCCC 1 | 525027 | 45.9 | 2.8 | 0.036 | 43 | 0.002 | 46 | 0.926 | 2.012 | Study
C6000-B:C2 | ASPARTATE TRANSAMINASE | s-PCCC 2-B | 461597 | 149 | 9 | 0.109 | 33 | 0.007 | 149.979 | 3.613 | 2.409 | Study
C6000-B:C2 | ASPARTATE TRANSAMINASE | PCCC 1-B | 525205 | 44.7 | 2.7 | 0.076 | 139 | 0.005 | 44.905 | 1.259 | 2.804 | In use
C6000-B:C2 | ASPARTATE TRANSAMINASE | PCCC 2-B | 535719 | 143 | 9 | -0.42 | 88 | -0.026 | 139.222 | 15.344 | 11.021 | In use
C6000-B:C2 | ASPARTATE TRANSAMINASE | PCCC 2-B | 519198 | 146 | 9 | 0.086 | 302 | 0.005 | 146.775 | 5.15 | 3.509 | Study
C6000-A:C1 | ASPARTATE TRANSAMINASE | PCCC 2 | 519198 | 146 | 9 | -0.261 | 344 | -0.016 | 143.651 | 3.103 | 2.16 | Study
C6000-B:C2 | ASPARTATE TRANSAMINASE | s-PCCC 1-B | 525027 | 45.9 | 2.8 | 0.164 | 38 | 0.01 | 46.361 | 1.047 | 2.259 | Study
C6000-A:C1 | ASPARTATE TRANSAMINASE | s-PCCC 2 | 461597 | 149 | 9 | -0.108 | 37 | -0.007 | 148.027 | 2.028 | 1.37 | Study
C6000-B:C2 | ASPARTATE TRANSAMINASE | PCCC 1-B | 525027 | 45.9 | 2.8 | 0.219 | 264 | 0.013 | 46.514 | 1.834 | 3.943 | Study
C6000-A:C1 | ASPARTATE TRANSAMINASE | PCCC 1 | 525027 | 45.9 | 2.8 | -0.143 | 307 | -0.009 | 45.498 | 1.101 | 2.42 | Study
C6000-A:E1-1 | BETA-HCG | RIAP L3 | 2110EC | 310 | 31 | -0.299 | 151 | -0.03 | 300.732 | 12.793 | 4.254 | Study
C6000-B:E3-2 | BETA-HCG | RIAP L3-B | 2147EC | 325 | 32.5 | -0.598 | 148 | -0.06 | 305.552 | 19.77 | 6.47 | Study
C6000-A:E1-1 | BETA-HCG | RIAP L3 | 2181EC | 272 | 27 | -0.714 | 151 | -0.071 | 252.719 | 12.844 | 5.082 | In use
C6000-B:E3-2 | BETA-HCG | RIAP L3-B | 2181EC | 272 | 27 | -0.513 | 154 | -0.051 | 258.142 | 16.922 | 6.555 | In use
C6000-B:E3-2 | BETA-HCG | RIAP L3-B | 2110EC | 310 | 31 | -0.51 | 144 | -0.051 | 294.176 | 13.432 | 4.566 | Study
C6000-A:E1-1 | BETA-HCG | RIAP L1 | 2179EC | 11.4 | 1.14 | -0.102 | 147 | -0.01 | 11.284 | 0.624 | 5.528 | In use
C6000-A:E1-1 | BETA-HCG | RIAP L1 | 2105EC | 11.1 | 0.83 | -0.096 | 324 | -0.007 | 11.02 | 0.514 | 4.668 | Study
C6000-B:E3-2 | BETA-HCG | RIAP L1-B | 2105EC | 11.1 | 0.833 | 0.041 | 292 | 0.003 | 11.134 | 0.884 | 7.94 | Study
C6000-B:E3-2 | BETA-HCG | RIAP L1-B | 2179EC | 11.4 | 1.14 | 0.06 | 141 | 0.006 | 11.468 | 0.681 | 5.935 | In use
C6000-A:E1-1 | BETA-HCG | RIAP L3 | 2147 | 325 | 32.5 | -0.907 | 165 | -0.091 | 295.538 | 12.347 | 4.178 | Study
C6000-B:C2 | BICARBONATE | AMM-P | 644006 | 30.5 | 1.8 | -0.748 | 94 | -0.044 | 29.154 | 1.425 | 4.887 | In use
C6000-B:C2 | BICARBONATE | AMM-P | 623128 | 30.7 | 1.8 | -0.567 | 65 | -0.033 | 29.68 | 1.129 | 3.804 | Study
C6000-B:C2 | BICARBONATE | AMM-N | 587880 | 18.8 | 1.1 | -0.022 | 219 | -0.001 | 18.775 | 0.931 | 4.959 | Study
C6000-B:C2 | BICARBONATE | AMM-P | 587881 | 33.8 | 2 | -0.818 | 214 | -0.048 | 32.164 | 1.089 | 3.387 | Study
C6000-B:C2 | BICARBONATE | AMM-N | 644004 | 17 | 1 | 0.956 | 160 | 0.056 | 17.956 | 0.888 | 4.943 | In use
C6000-A:E2-2 | C-peptide | RIAP L2 | 2107EC | 6.22 | 0.77 | 0.06 | 357 | 0.007 | 6.266 | 0.22 | 3.51 | Study
C6000-A:E2-2 | C-peptide | RIAP L1 | 2105EC | 2.1 | 0.265 | 0.013 | 382 | 0.002 | 2.103 | 0.091 | 4.315 | Study
C6000-A:E2-2 | C-peptide | RIAP L1 | 2179EC | 1.94 | 0.16 | -1.025 | 185 | -0.085 | 1.776 | 0.129 | 7.279 | In use
C6000-A:E2-2 | C-peptide | RIAP L2 | 2124EC | 7.16 | 0.785 | -0.606 | 105 | -0.066 | 6.684 | 0.313 | 4.689 | Study
C6000-A:E2-2 | C-peptide | RIAP L2 | 2228EC | 6.16 | 0.77 | -0.12 | 89 | -0.015 | 6.068 | 0.224 | 3.692 | In use
C6000-A:C1 | C-REACTIVE PROTEIN | s-PCCC 1 | 525027 | 9.05 | 0.6 | -0.408 | 48 | -0.027 | 8.805 | 0.241 | 2.74 | Study
C6000-B:C2 | C-REACTIVE PROTEIN | PCCC 2-B | 519198 | 50.89 | 3.6 | 0.123 | 316 | 0.009 | 51.332 | 1.669 | 3.251 | Study
C6000-A:C1 | C-REACTIVE PROTEIN | PCCC 2 | 519198 | 50.89 | 3.6 | 0.015 | 366 | 0.001 | 50.944 | 1.66 | 3.259 | Study
C6000-A:C1 | C-REACTIVE PROTEIN | PCCC 1 | 525027 | 8.76 | 0.6 | 0.082 | 321 | 0.006 | 8.809 | 0.286 | 3.247 | Study
C6000-B:C2 | C-REACTIVE PROTEIN | s-PCCC 1-B | 525027 | 9.05 | 0.6 | -0.532 | 41 | -0.035 | 8.731 | 0.181 | 2.074 | Study
C6000-B:C2 | C-REACTIVE PROTEIN | PCCC 1-B | 525027 | 8.76 | 0.6 | 0.009 | 261 | 0.001 | 8.765 | 0.355 | 4.046 | Study
C6000-B:C2 | C-REACTIVE PROTEIN | PCCC 1-B | 525205 | 5.68 | 0.21 | -0.716 | 139 | -0.026 | 5.53 | 0.542 | 9.803 | In use
C6000-A:C1 | C-REACTIVE PROTEIN | s-PCCC 2 | 461597 | 52.7 | 3.5 | -0.174 | 43 | -0.012 | 52.091 | 2.449 | 4.701 | Study
C6000-B:C2 | C-REACTIVE PROTEIN | s-PCCC 2-B | 461597 | 54.3 | 3.6 | -0.01 | 37 | -0.001 | 54.266 | 1.383 | 2.549 | Study
C6000-B:C2 | C-REACTIVE PROTEIN | PCCC 2-B | 535719 | 51.4 | 3.4 | -0.131 | 85 | -0.009 | 50.955 | 5.763 | 11.31 | In use
C6000-A:C1 | C-REACTIVE PROTEIN | PCCC 2 | 535719 | 51.4 | 3.4 | -0.012 | 98 | -0.001 | 51.359 | 1.39 | 2.707 | In use
C6000-A:C1 | C-REACTIVE PROTEIN | PCCC 1 | 525205 | 5.68 | 0.21 | -0.079 | 156 | -0.003 | 5.663 | 0.129 | 2.276 | In use
C6000-A:E2-1 | CA 125 | RIAP L1 | 2179EC | 19 | 1.9 | 0.585 | 121 | 0.058 | 20.111 | 0.777 | 3.863 | In use
C6000-A:E2-1 | CA 125 | RIAP L3 | 2147 | 187 | 18.5 | 0.209 | 155 | 0.021 | 190.874 | 10.549 | 5.527 | Study
C6000-A:E2-1 | CA 125 | RIAP L1 | 2105EC | 18.6 | 1.65 | -0.057 | 302 | -0.005 | 18.505 | 0.856 | 4.625 | Study
C6000-A:E2-1 | CA 125 | RIAP L3 | 2110EC | 183 | 18.5 | 0.257 | 139 | 0.026 | 187.761 | 6.109 | 3.253 | Study
C6000-A:E2-1 | CA 125 | RIAP L3 | 2181EC | 182 | 18 | 0.761 | 133 | 0.075 | 195.694 | 4.912 | 2.51 | In use
C6000-A:E2-1 | CA 19-9 | RIAP L3 | 2110EC | 168.45 | 16 | -0.022 | 146 | -0.002 | 168.101 | 6.475 | 3.852 | Study
C6000-A:E2-1 | CA 19-9 | RIAP L1 | 2179EC | 20.8 | 2.1 | -0.401 | 124 | -0.04 | 19.958 | 1.151 | 5.767 | In use
C6000-A:E2-1 | CA 19-9 | RIAP L1 | 2105EC | 23.9 | 1.79 | -0.169 | 327 | -0.013 | 23.598 | 1.55 | 6.566 | Study
C6000-A:E2-1 | CA 19-9 | RIAP L3 | 2181EC | 165 | 16.5 | 0.1 | 138 | 0.01 | 166.654 | 7.163 | 4.298 | In use
C6000-A:E2-1 | CA 19-9 | RIAP L3 | 2147 | 168 | 17 | -0.297 | 164 | -0.03 | 162.957 | 9.048 | 5.553 | Study
C6000-B:C2 | CALCIUM | PCCC 2-B | 519198 | 3.38 | 0.14 | -0.214 | 371 | -0.009 | 3.35 | 0.193 | 5.748 | Study
C6000-B:C2 | CALCIUM | PCCC 2-B | 535719 | 3.47 | 0.14 | -0.211 | 96 | -0.008 | 3.441 | 0.036 | 1.037 | In use
C6000-A:C1 | CALCIUM | PCCC 2 | 519198 | 3.38 | 0.14 | 0.106 | 378 | 0.004 | 3.395 | 0.066 | 1.949 | Study
C6000-B:C2 | CALCIUM | PCCC 2-B | 461597 | 3.4 | 0.06 | -0.056 | 74 | -0.001 | 3.397 | 0.06 | 1.752 | Study
C6000-A:C1 | CALCIUM | PCCC 2 | 535719 | 3.47 | 0.14 | -0.088 | 94 | -0.004 | 3.458 | 0.05 | 1.453 | In use
C6000-A:C1 | CALCIUM | PCCC 2 | 461597 | 3.4 | 0.06 | -0.15 | 70 | -0.003 | 3.391 | 0.031 | 0.904 | Study
C6000-B:C2 | CALCIUM | PCCC 1-B | 525205 | 2.14 | 0.09 | -0.258 | 156 | -0.011 | 2.117 | 0.046 | 2.176 | In use
C6000-A:C1 | CALCIUM | PCCC 1 | 525205 | 2.14 | 0.09 | 0.067 | 148 | 0.003 | 2.146 | 0.043 | 2.009 | In use
C6000-B:C2 | CALCIUM | PCCC 1-B | 525027 | 2.26 | 0.09 | -0.439 | 393 | -0.017 | 2.22 | 0.046 | 2.089 | Study
C6000-A:C1 | CALCIUM | PCCC 1 | 525027 | 2.24 | 0.09 | -0.032 | 409 | -0.001 | 2.237 | 0.04 | 1.801 | Study
C6000-B:C2 | CALCIUM URINE | Ran-Liq 3 | 1238UC | 2.78 | 0.14 | 0.564 | 127 | 0.028 | 2.859 | 0.082 | 2.866 | In use
C6000-B:C2 | CALCIUM URINE | Ran-Liq 2 | 1209UC | 1.95 | 0.1 | 0.288 | 224 | 0.015 | 1.979 | 0.044 | 2.219 | Study
C6000-B:C2 | CALCIUM URINE | Ran-Liq 3 | 1210UC | 2.83 | 0.14 | 0.476 | 224 | 0.024 | 2.897 | 0.073 | 2.505 | Study
C6000-B:C2 | CALCIUM URINE | Bio-liq1 | 88161 | 1.79 | 0.147 | 0.244 | 122 | 0.02 | 1.826 | 0.045 | 2.447 | Study
C6000-B:C2 | CALCIUM URINE | Bio-liq 2 | 88162 | 2.64 | 0.09 | 0.777 | 112 | 0.026 | 2.71 | 0.047 | 1.74 | Study
C6000-B:C2 | CALCIUM URINE | Ran-Liq 2 | 1237UC | 1.87 | 0.1 | 0.551 | 116 | 0.029 | 1.925 | 0.059 | 3.045 | In use
C6000-A:E1-2 | CARCINOEMBRYONIC AG | RIAP L1 | 2105EC | 4.18 | 0.26 | 0.766 | 314 | 0.048 | 4.379 | 0.174 | 3.964 | Study
C6000-A:E1-2 | CARCINOEMBRYONIC AG | RIAP L1 | 2179EC | 4.08 | 0.41 | 0.301 | 120 | 0.03 | 4.204 | 0.258 | 6.127 | In use
C6000-A:E1-2 | CARCINOEMBRYONIC AG | RIAP L3 | 2181EC | 50 | 5 | 0.174 | 134 | 0.017 | 50.868 | 4.181 | 8.219 | In use
C6000-A:E1-2 | CARCINOEMBRYONIC AG | RIAP L3 | 2147 | 52.2 | 5.2 | 0.13 | 151 | 0.013 | 52.876 | 3.542 | 6.699 | Study
C6000-A:E1-2 | CARCINOEMBRYONIC AG | RIAP L3 | 2110EC | 51.6 | 5.15 | 0.416 | 143 | 0.042 | 53.743 | 2.253 | 4.191 | Study
C6000-B:ISE1 | CHLORIDE | PCCC 1-B | 525027 | 87.3 | 2.7 | 0.33 | 368 | 0.01 | 88.191 | 1.554 | 1.762 | Study
C6000-B:ISE1 | CHLORIDE | PCCC 2-B | 519198 | 106 | 3 | -0.053 | 278 | -0.002 | 105.84 | 1.586 | 1.499 | Study
C6000-B:ISE1 | CHLORIDE | PCCC 2-B | 461597 | 103.1 | 3 | 0.157 | 92 | 0.005 | 103.572 | 1.375 | 1.328 | Study
C6000-A:ISE1 | CHLORIDE | PCCC 1 | 525027 | 87.3 | 2.7 | -0.07 | 106 | -0.002 | 87.111 | 1.403 | 1.611 | Study
C6000-B:ISE1 | CHLORIDE | s-PCCC 2-B | 535719 | 106 | 3 | 0.075 | 188 | 0.002 | 106.224 | 3.564 | 3.356 | In use
C6000-B:ISE1 | CHLORIDE | s-PCCC 1-B | 525027 | 87.3 | 2.7 | 0.414 | 565 | 0.013 | 88.418 | 1.567 | 1.773 | Study
C6000-A:ISE1 | CHLORIDE | PCCC 2 | 461597 | 103.1 | 3 | 0.044 | 88 | 0.001 | 103.232 | 1.804 | 1.747 | Study
C6000-B:ISE1 | CHLORIDE | s-PCCC 1-B | 525205 | 83.7 | 2.5 | 0.912 | 303 | 0.027 | 85.98 | 9.101 | 10.585 | In use
C6000-A:ISE1 | CHLORIDE | s-PCCC 1 | 525205 | 83.7 | 2.5 | 0.298 | 342 | 0.009 | 84.446 | 1.836 | 2.174 | In use
C6000-A:ISE1 | CHLORIDE | s-PCCC 2 | 519198 | 103 | 3 | 0.745 | 712 | 0.022 | 105.235 | 1.839 | 1.747 | Study
C6000-A:ISE1 | CHLORIDE | s-PCCC 1 | 525027 | 87.3 | 2.7 | 0.094 | 595 | 0.003 | 87.555 | 1.913 | 2.184 | Study
C6000-A:ISE1 | CHLORIDE | s-PCCC 2 | 535719 | 103 | 3 | 0.816 | 193 | 0.024 | 105.448 | 1.748 | 1.658 | In use
C6000-B:ISE1 | CHLORIDE | s-PCCC 2-B | 519198 | 106 | 3 | 0.02 | 696 | 0.001 | 106.061 | 1.948 | 1.837 | Study
C6000-B:ISE1 | CHLORIDE URINE | Bio-liq1 | 88161 | 89.6 | 6.15 | 0.702 | 165 | 0.048 | 93.919 | 2.093 | 2.229 | Study
C6000-B:ISE1 | CHLORIDE URINE | Bio-liq 2 | 88162 | 188 | 8 | 0.773 | 157 | 0.033 | 194.188 | 3.983 | 2.051 | Study
C6000-B:ISE1 | CHLORIDE URINE | Ran-Liq 3 | 1238UC | 181 | 13.5 | -0.041 | 154 | -0.003 | 180.447 | 4.717 | 2.614 | In use
C6000-B:ISE1 | CHLORIDE URINE | Ran-Liq 3 | 1210UC | 170 | 12.5 | -0.148 | 252 | -0.011 | 168.146 | 6.517 | 3.876 | Study
C6000-B:ISE1 | CHLORIDE URINE | Ran-Liq 2 | 1237UC | 76.2 | 5.7 | 0.593 | 138 | 0.044 | 79.582 | 1.68 | 2.111 | In use
C6000-B:ISE1 | CHLORIDE URINE | Ran-Liq 2 | 1209UC | 69.6 | 5.2 | -0.269 | 246 | -0.02 | 68.199 | 3.002 | 4.402 | Study
C6000-B:C2 | COMPLEMENT C3 | PCCC 2-B | 519198 | 1.57 | 0.13 | 0.142 | 248 | 0.012 | 1.588 | 0.062 | 3.932 | Study
C6000-B:C2 | COMPLEMENT C3 | s-PCCC 1-B | 525027 | 0.933 | 0.075 | 0.371 | 30 | 0.03 | 0.961 | 0.015 | 1.612 | Study
C6000-B:C2 | COMPLEMENT C3 | PCCC 1-B | 525027 | 0.933 | 0.075 | 0.541 | 212 | 0.043 | 0.974 | 0.03 | 3.035 | Study
C6000-B:C2 | COMPLEMENT C3 | PCCC 1-B | 525205 | 0.926 | 0.074 | 0.165 | 109 | 0.013 | 0.938 | 0.039 | 4.16 | In use
C6000-B:C2 | COMPLEMENT C3 | PCCC 2-B | 535719 | 1.55 | 0.12 | -0.361 | 65 | -0.028 | 1.507 | 0.07 | 4.66 | In use
C6000-B:C2 | COMPLEMENT C4 | PCCC 1-B | 525205 | 0.145 | 0.012 | -0.508 | 109 | -0.042 | 0.139 | 0.003 | 2.196 | In use
C6000-B:C2 | COMPLEMENT C4 | PCCC 1-B | 525027 | 0.13 | 0.01 | -0.566 | 195 | -0.044 | 0.124 | 0.004 | 3.519 | Study
C6000-B:C2 | COMPLEMENT C4 | PCCC 2-B | 519198 | 0.224 | 0.018 | -0.315 | 231 | -0.025 | 0.218 | 0.007 | 3.09 | Study
C6000-B:C2 | COMPLEMENT C4 | PCCC 2-B | 535719 | 0.227 | 0.018 | -0.197 | 66 | -0.016 | 0.223 | 0.003 | 1.484 | In use
C6000-A:C1 | CONJUGATED BILIRUBIN | PCCC 2 | 461597 | 43.7 | 3.5 | -0.73 | 69 | -0.058 | 41.146 | 1.848 | 4.491 | Study
C6000-B:C2 | CONJUGATED BILIRUBIN | PCCC 1-B | 525027 | 15.6 | 1 | -0.126 | 354 | -0.008 | 15.474 | 0.733 | 4.739 | Study
C6000-B:C2 | CONJUGATED BILIRUBIN | PCCC 2-B | 535719 | 44.9 | 3.6 | -0.389 | 93 | -0.031 | 43.499 | 1.56 | 3.587 | In use
C6000-A:C1 | CONJUGATED BILIRUBIN | PCCC 1 | 525027 | 15.2 | 1.3 | -0.198 | 393 | -0.017 | 14.943 | 0.954 | 6.381 | Study
C6000-B:C2 | CONJUGATED BILIRUBIN | PCCC 2-B | 519198 | 41.9 | 3.5 | 0.208 | 328 | 0.017 | 42.629 | 2.105 | 4.937 | Study
C6000-B:C2 | CONJUGATED BILIRUBIN | PCCC 1-B | 525205 | 16.4 | 1.3 | -0.183 | 148 | -0.015 | 16.161 | 0.611 | 3.781 | In use
C6000-A:C1 | CONJUGATED BILIRUBIN | MAS BILIRUBIN 3 | BC24033A | 177 | 17.3 | -0.244 | 386 | -0.024 | 172.779 | 6.978 | 4.039 | In use
C6000-B:C2 | CONJUGATED BILIRUBIN | MAS BILIRUBIN-B | BC24033A | 177 | 17.3 | -0.129 | 338 | -0.013 | 174.774 | 7.144 | 4.088 | In use
C6000-B:C2 | CONJUGATED BILIRUBIN | PCCC 2-B | 461597 | 43.7 | 3.5 | -0.568 | 67 | -0.046 | 41.71 | 0.94 | 2.253 | Study
C6000-A:C1 | CONJUGATED BILIRUBIN | PCCC 1 | 525205 | 16.4 | 1.3 | -0.159 | 151 | -0.013 | 16.193 | 0.575 | 3.551 | In use
C6000-A:C1 | CONJUGATED BILIRUBIN | PCCC 2 | 519198 | 40.7 | 3.5 | 0.304 | 354 | 0.026 | 41.764 | 1.786 | 4.276 | Study
C6000-A:C1 | CONJUGATED BILIRUBIN | PCCC 2 | 535719 | 44.9 | 3.6 | -0.236 | 98 | -0.019 | 44.05 | 1.319 | 2.994 | In use
C6000-A:C1 | CONJUGATED BILIRUBIN | MAS BILIRUBIN 3 | BC23063A | 173 | 17.3 | -0.343 | 66 | -0.034 | 167.064 | 7.828 | 4.686 | Study
C6000-B:C2 | CONJUGATED BILIRUBIN | MAS BILIRUBIN-B | BC23063A | 173 | 17.3 | -0.163 | 61 | -0.016 | 170.185 | 4.929 | 2.896 | Study
C6000-A:E1-2 | CORTISOL II | RIAP L1 | 2105EC | 196 | 11.17 | 0.129 | 300 | 0.007 | 197.438 | 8.182 | 4.144 | Study
C6000-A:E1-2 | CORTISOL II | RIAP L2 | 2228EC | 660 | 82.5 | 0.421 | 63 | 0.053 | 694.73 | 42.703 | 6.147 | In use
C6000-A:E1-2 | CORTISOL II | RIAP L1 | 2179EC | 201 | 25 | -0.083 | 117 | -0.01 | 198.93 | 11.613 | 5.838 | In use
C6000-A:E1-2 | CORTISOL II | RIAP L2 | 2124EC | 638 | 77 | 0.04 | 74 | 0.005 | 641.069 | 36.783 | 5.738 | Study
C6000-A:E1-2 | CORTISOL II | RIAP L2 | 2107EC | 631 | 79 | 0.158 | 281 | 0.02 | 643.459 | 24.39 | 3.79 | Study
C6000-A:C1 | CREATINE KINASE Gen 2 | s-PCCC 2 | 461597 | 266 | 16 | -0.411 | 36 | -0.025 | 259.417 | 3.549 | 1.368 | Study
C6000-A:C1 | CREATINE KINASE Gen 2 | PCCC 1 | 525027 | 154 | 9 | -0.499 | 289 | -0.029 | 149.512 | 3.228 | 2.159 | Study
C6000-A:C1 | CREATINE KINASE Gen 2 | s-PCCC 1 | 525027 | 154 | 9 | -0.485 | 44 | -0.028 | 149.636 | 1.672 | 1.117 | Study
C6000-A:C1 | CREATINE KINASE Gen 2 | PCCC 1 | 525205 | 153 | 9 | -0.531 | 134 | -0.031 | 148.224 | 2.076 | 1.401 | In use
C6000-A:C1 | CREATINE KINASE Gen 2 | PCCC 2 | 535719 | 267 | 16 | -0.368 | 82 | -0.022 | 261.11 | 3.768 | 1.443 | In use
C6000-A:C1 | CREATINE KINASE Gen 2 | PCCC 2 | 519198 | 264 | 16 | -0.469 | 329 | -0.028 | 256.502 | 4.949 | 1.93 | Study
C6000-A:C1 | CREATININE | PCCC 2 | 461597 | 355 | 21 | 0.294 | 136 | 0.017 | 361.184 | 5.952 | 1.648 | Study
C6000-A:C1 | CREATININE | PCCC 1 | 525205 | 88.7 | 5.3 | 0.118 | 417 | 0.007 | 89.326 | 1.994 | 2.232 | In use
C6000-B:C2 | CREATININE | PCCC 1-B | 525027 | 87.6 | 5.3 | 0.26 | 762 | 0.016 | 88.979 | 2.124 | 2.387 | Study
C6000-A:C1 | CREATININE | PCCC 1 | 525027 | 88.5 | 5.3 | 0.169 | 979 | 0.01 | 89.396 | 2.049 | 2.292 | Study
C6000-B:C2 | CREATININE | PCCC 2-B | 519198 | 347 | 21 | 0.016 | 759 | 0.001 | 347.335 | 7.867 | 2.265 | Study
C6000-B:C2 | CREATININE | PCCC 2-B | 461597 | 355 | 21 | -0.147 | 121 | -0.009 | 351.909 | 8.885 | 2.525 | Study
C6000-A:C1 | CREATININE | PCCC 2 | 535719 | 329 | 20 | 0.097 | 259 | 0.006 | 330.95 | 7.051 | 2.131 | In use
C6000-A:C1 | CREATININE | PCCC 2 | 519198 | 347 | 21 | 0.252 | 982 | 0.015 | 352.287 | 6.523 | 1.852 | Study
C6000-B:C2 | CREATININE | PCCC 1-B | 525205 | 88.7 | 5.3 | -0.223 | 359 | -0.013 | 87.518 | 1.62 | 1.851 | In use
C6000-B:C2 | CREATININE | PCCC 2-B | 535719 | 329 | 20 | -0.3 | 231 | -0.018 | 322.996 | 4.979 | 1.541 | In use
C6000-A:C1 | CREATININE URINE | Ran-Liq 3 | 1210UC | 15.6 | 1.55 | -0.704 | 119 | -0.07 | 14.508 | 0.545 | 3.757 | Study
C6000-B:C2 | CREATININE URINE | Ran-Liq 3 | 1238UC | 15.5 | 1.55 | -0.23 | 33 | -0.023 | 15.144 | 0.181 | 1.196 | In use
C6000-B:C2 | CREATININE URINE | Ran-Liq 2 | 1237UC | 6.68 | 0.67 | -0.214 | 145 | -0.022 | 6.536 | 0.139 | 2.126 | In use
C6000-B:C2 | CREATININE URINE | Ran-Liq 2 | 1209UC | 6.86 | 0.69 | -0.622 | 299 | -0.063 | 6.431 | 0.168 | 2.609 | Study
C6000-A:C1 | CREATININE URINE | Ran-Liq 2 | 1237UC | 6.68 | 0.67 | -0.299 | 114 | -0.03 | 6.48 | 0.191 | 2.953 | In use
C6000-A:C1 | CREATININE URINE | PP PUC | 64692700 | 4.71 | 0.28 | -0.542 | 186 | -0.032 | 4.558 | 0.097 | 2.128 | Study
C6000-A:C1 | CREATININE URINE | PN PUC | 649811 | 9.11 | 0.55 | -0.485 | 401 | -0.029 | 8.843 | 0.183 | 2.072 | Study
C6000-A:C1 | CREATININE URINE | Ran-Liq 2 | 1209UC | 6.86 | 0.69 | -0.729 | 124 | -0.073 | 6.357 | 0.213 | 3.352 | Study
C6000-A:C1 | CREATININE URINE | PN PUC | 602712 | 8.67 | 0.52 | -0.28 | 295 | -0.017 | 8.525 | 0.234 | 2.746 | Study
C6000-A:C1 | CREATININE URINE | PN PUC | 531179 | 9.05 | 0.54 | -0.799 | 174 | -0.048 | 8.619 | 0.328 | 3.803 | Study
C6000-A:C1 | CREATININE URINE | PN PUC | 634291 | 9.09 | 0.55 | -0.308 | 37 | -0.019 | 8.92 | 0.122 | 1.365 | Study
C6000-A:C1 | CREATININE URINE | PP PUC | 602710 | 4.66 | 0.28 | -0.46 | 385 | -0.028 | 4.531 | 0.13 | 2.867 | Study
C6000-A:C1 | CREATININE URINE | PP PUC | 531181 | 4.35 | 0.26 | 0.387 | 80 | 0.023 | 4.451 | 0.165 | 3.702 | Study
C6000-A:C1 | CREATININE URINE | Ran-Liq 3 | 1238UC | 15.5 | 1.55 | -0.478 | 112 | -0.048 | 14.759 | 0.386 | 2.615 | In use
C6000-A:C1 | CREATININE URINE | PP PUC | 637091 | 4.51 | 0.27 | 0.045 | 239 | 0.003 | 4.522 | 0.096 | 2.125 | Study
C6000-A:E2-2 | DEHYDROEPIANDROSTERONE SULPHAT | RIAP L2 | 2107EC | 16.7 | 1.09 | -0.66 | 277 | -0.043 | 15.98 | 0.774 | 4.841 | Study
C6000-A:E2-2 | DEHYDROEPIANDROSTERONE SULPHAT | RIAP L1 | 2179EC | 2.68 | 0.27 | -0.167 | 117 | -0.017 | 2.635 | 0.179 | 6.794 | In use
C6000-A:E2-2 | DEHYDROEPIANDROSTERONE SULPHAT | RIAP L2 | 2124EC | 17.2 | 1.12 | -0.312 | 80 | -0.02 | 16.85 | 0.838 | 4.971 | Study
C6000-A:E2-2 | DEHYDROEPIANDROSTERONE SULPHAT | RIAP L1 | 2105EC | 2.81 | 0.184 | -0.083 | 291 | -0.005 | 2.795 | 0.151 | 5.394 | Study
C6000-A:E2-2 | DEHYDROEPIANDROSTERONE SULPHAT | RIAP L2 | 2228EC | 17.5 | 1.75 | -0.103 | 68 | -0.01 | 17.321 | 1.239 | 7.152 | In use
C6000-A:E2-1 | FERRITIN | RIAP L3 | 2181EC | 318 | 32 | 0.714 | 163 | 0.072 | 340.861 | 14.466 | 4.244 | In use
C6000-A:E2-1 | FERRITIN | sRIAP L1 | 2179EC | 20.57 | 2.03 | 1.097 | 290 | 0.108 | 22.798 | 1.591 | 6.977 | In use
C6000-A:E2-1 | FERRITIN | PC V1 | 595763 | 146 | 13.1 | -0.119 | 44 | -0.011 | 144.439 | 2.954 | 2.045 | Study
C6000-A:E2-1 | FERRITIN | sRIAP L3 | 2147 EC | 300 | 30 | 0.391 | 325 | 0.039 | 311.743 | 11.698 | 3.752 | Study
C6000-A:E2-1 | FERRITIN | sRIAP L3 | 2181EC | 318 | 32 | 0.662 | 307 | 0.067 | 339.172 | 12.654 | 3.731 | In use
C6000-A:E2-1 | FERRITIN | sRIAP L1 | 2105EC | 14.05 | 1.18 | 0.237 | 515 | 0.02 | 14.33 | 1.274 | 8.887 | Study
C6000-A:E2-1 | FERRITIN | RIAP L1 | 2105EC | 15.7 | 1.35 | 0.008 | 133 | 0.001 | 15.711 | 0.932 | 5.935 | Study
C6000-A:E2-1 | FERRITIN | RIAP L3 | 2110EC | 310 | 31 | 0.536 | 133 | 0.054 | 326.623 | 12.159 | 3.723 | Study
C6000-A:E1-1 | FOLLICLE STIMULATING HORMONE | RIAP L1 | 2105EC | 6.62 | 0.35 | 0.198 | 297 | 0.01 | 6.689 | 0.216 | 3.225 | Study
C6000-A:E1-2 | FOLLICLE STIMULATING HORMONE | RIAP L3 | 2181EC | 54.2 | 5.4 | 0.42 | 61 | 0.042 | 56.469 | 1.421 | 2.516 | In use
C6000-A:E1-1 | FOLLICLE STIMULATING HORMONE | RIAP L1 | 2179EC | 6.03 | 0.605 | -0.092 | 68 | -0.009 | 5.974 | 0.208 | 3.474 | In use
C6000-A:E1-2 | FOLLICLE STIMULATING HORMONE | RIAP L1 | 2179EC | 6.03 | 0.605 | 0.781 | 62 | 0.078 | 6.502 | 0.185 | 2.846 | In use
C6000-A:E1-1 | FOLLICLE STIMULATING HORMONE | RIAP L3 | 2147 | 55.3 | 5.55 | -0.227 | 148 | -0.023 | 54.04 | 1.482 | 2.743 | Study
C6000-A:E1-1 | FOLLICLE STIMULATING HORMONE | RIAP L3 | 2181EC | 54.2 | 5.4 | -0.319 | 71 | -0.032 | 52.477 | 1.457 | 2.777 | In use
C6000-A:E1-1 | FOLLICLE STIMULATING HORMONE | RIAP L3 | 2110EC | 55.3 | 5.55 | 0.269 | 140 | 0.027 | 56.794 | 2.46 | 4.331 | Study
C6000-A:E1-2 | FREE PSA | RIAP L1 | 2179EC | 1.19 | 0.149 | -0.245 | 114 | -0.031 | 1.154 | 0.045 | 3.884 | In use
C6000-A:E1-2 | FREE PSA | RIAP L2 | 2228EC | 9.21 | 1.15 | -0.052 | 60 | -0.006 | 9.151 | 0.234 | 2.554 | In use
C6000-A:E1-2 | FREE PSA | RIAP L2 | 2107EC | 7.46 | 0.933 | 0.275 | 290 | 0.034 | 7.716 | 0.43 | 5.567 | Study
C6000-A:E1-2 | FREE PSA | RIAP L2 | 2124EC | 9.78 | 0.895 | 0.251 | 75 | 0.023 | 10.005 | 0.286 | 2.861 | Study
C6000-A:E1-2 | FREE PSA | RIAP L1 | 2105EC | 1.22 | 0.092 | 0.051 | 302 | 0.004 | 1.225 | 0.064 | 5.266 | Study
C6000-A:E1-1 | Free T4 Gen 3 | sRIAP L2 | 2228EC | 29.9 | 3.75 | -0.231 | 126 | -0.029 | 29.033 | 1.217 | 4.191 | In use
C6000-A:E1-1 | Free T4 Gen 3 | RIAP L2 | 2107EC | 29.1 | 3.65 | -0.142 | 144 | -0.018 | 28.582 | 1.329 | 4.649 | Study
C6000-A:E1-1 | Free T4 Gen 3 | sRIAP L1 | 2105EC | 13.3 | 0.53 | 0.373 | 476 | 0.015 | 13.497 | 0.504 | 3.737 | Study
C6000-A:E1-1 | Free T4 Gen 3 | sRIAP L2 | 2107EC | 29.1 | 3.65 | 0.004 | 434 | 0.001 | 29.116 | 1.167 | 4.008 | Study
C6000-A:E1-1 | Free T4 Gen 3 | sRIAP L2 | 2124EC | 29.5 | 1.168 | 0.137 | 147 | 0.005 | 29.66 | 0.736 | 2.48 | Study
C6000-A:E1-1 | Free T4 Gen 3 | sRIAP L1 | 2179EC | 13.7 | 1.7 | -0.06 | 250 | -0.007 | 13.597 | 0.572 | 4.209 | In use
C6000-A:E1-1 | Free T4 Gen 3 | RIAP L1 | 2105EC | 13.3 | 0.53 | -0.407 | 131 | -0.016 | 13.084 | 0.47 | 3.59 | Study
C6000-B:C2 | GAMMA-GLUTAMYL TRANSFERASE | s-PCCC 2-B | 461597 | 278 | 17 | -0.228 | 40 | -0.014 | 274.125 | 2.244 | 0.819 | Study
C6000-A:C1 | GAMMA-GLUTAMYL TRANSFERASE | PCCC 2 | 519198 | 225 | 14 | -0.29 | 349 | -0.018 | 220.934 | 3.569 | 1.616 | Study
C6000-A:C1 | GAMMA-GLUTAMYL TRANSFERASE | s-PCCC 1 | 525027 | 55.8 | 3.3 | -0.178 | 52 | -0.011 | 55.212 | 1.903 | 3.447 | Study
C6000-A:C1 | GAMMA-GLUTAMYL TRANSFERASE | s-PCCC 2 | 461597 | 278 | 17 | -0.408 | 45 | -0.025 | 271.067 | 10.226 | 3.772 | Study
C6000-B:C2 | GAMMA-GLUTAMYL TRANSFERASE | PCCC 2-B | 519198 | 225 | 14 | -0.256 | 337 | -0.016 | 221.421 | 6.641 | 2.999 | Study
C6000-B:C2 | GAMMA-GLUTAMYL TRANSFERASE | PCCC 1-B | 525205 | 58.2 | 3.5 | -0.339 | 144 | -0.02 | 57.014 | 1.194 | 2.095 | In use
C6000-A:C1 | GAMMA-GLUTAMYL TRANSFERASE | PCCC 1 | 525027 | 55.8 | 3.3 | -0.212 | 311 | -0.013 | 55.1 | 1.15 | 2.087 | Study
C6000-B:C2 | GAMMA-GLUTAMYL TRANSFERASE | s-PCCC 1-B | 525027 | 55.8 | 3.3 | -0.081 | 47 | -0.005 | 55.532 | 0.776 | 1.397 | Study
C6000-A:C1 | GAMMA-GLUTAMYL TRANSFERASE | PCCC 1 | 525205 | 58.2 | 3.5 | -0.226 | 142 | -0.014 | 57.408 | 1.198 | 2.087 | In use
C6000-B:C2 | GAMMA-GLUTAMYL TRANSFERASE | PCCC 1-B | 525027 | 55.8 | 3.3 | -0.308 | 291 | -0.018 | 54.784 | 1.62 | 2.958 | Study
C6000-A:C1 | GAMMA-GLUTAMYL TRANSFERASE | PCCC 2 | 535719 | 229 | 14 | -0.164 | 90 | -0.01 | 226.7 | 2.717 | 1.198 | In use
C6000-B:C2 | GAMMA-GLUTAMYL TRANSFERASE | PCCC 2-B | 535719 | 229 | 14 | -0.315 | 88 | -0.019 | 224.591 | 3.386 | 1.507 | In use
C6000-B:C2 | GLUCOSE | PCCC 1-B | 525205 | 5.76 | 0.29 | 0.031 | 132 | 0.002 | 5.769 | 0.169 | 2.927 | In use
C6000-A:C1 | GLUCOSE | PCCC 1 | 525205 | 5.76 | 0.29 | 0.093 | 130 | 0.005 | 5.787 | 0.127 | 2.197 | In use
C6000-A:C1 | GLUCOSE | PCCC 2 | 535719 | 13.6 | 0.7 | -0.215 | 82 | -0.011 | 13.45 | 0.305 | 2.27 | In use
C6000-B:C2 | GLUCOSE | PCCC 2-B | 519198 | 13.4 | 0.7 | -0.009 | 311 | 0 | 13.393 | 0.343 | 2.562 | Study
C6000-B:C2 | GLUCOSE | PCCC 2-B | 461597 | 13.2 | 0.7 | -0.323 | 62 | -0.017 | 12.974 | 0.204 | 1.572 | Study
C6000-A:C1 | GLUCOSE | PCCC 2 | 519198 | 13.4 | 0.7 | 0.1 | 326 | 0.005 | 13.47 | 0.277 | 2.056 | Study
C6000-A:C1 | GLUCOSE | PCCC 2 | 461597 | 13.2 | 0.7 | -0.143 | 66 | -0.008 | 13.1 | 0.205 | 1.562 | Study
C6000-B:C2 | GLUCOSE | PCCC 2-B | 535719 | 13.6 | 0.7 | -0.462 | 81 | -0.024 | 13.277 | 0.284 | 2.14 | In use
C6000-B:C2 | GLUCOSE | PCCC 1-B | 525027 | 5.74 | 0.29 | -0.134 | 337 | -0.007 | 5.701 | 0.121 | 2.121 | Study
C6000-A:C1 | GLUCOSE | PCCC 1 | 525027 | 5.74 | 0.29 | 0.067 | 360 | 0.003 | 5.759 | 0.118 | 2.048 | Study
C6000-A:C1 | GLYCATED HAEMOGLOBIN (NGSP) | PCA1N-VER | 617428 | 5.7 | 0.34 | 0.11 | 139 | 0.007 | 5.737 | 0.118 | 2.059 | In use
C6000-B:C2 | GLYCATED HAEMOGLOBIN (NGSP) | PCA1N | 563450 | 5.69 | 0.17 | 0.275 | 63 | 0.008 | 5.737 | 0.07 | 1.213 | Study
C6000-B:C2 | GLYCATED HAEMOGLOBIN (NGSP) | PCA1P | 563451 | 10.6 | 0.32 | 0.323 | 589 | 0.01 | 10.703 | 0.161 | 1.508 | Study
C6000-B:C2 | GLYCATED HAEMOGLOBIN (NGSP) | PCA1P | 617429 | 9.69 | 0.58 | 0.081 | 245 | 0.005 | 9.737 | 0.171 | 1.76 | In use
C6000-B:C2 | GLYCATED HAEMOGLOBIN (NGSP) | PCA1N | 509111 | 5.65 | 0.17 | 0.65 | 110 | 0.02 | 5.76 | 0.074 | 1.29 | Study
C6000-A:C1 | GLYCATED HAEMOGLOBIN (NGSP) | PCA1P | 617429 | 9.69 | 0.58 | -0.076 | 91 | -0.005 | 9.646 | 0.165 | 1.709 | In use
C6000-A:C1 | GLYCATED HAEMOGLOBIN (NGSP) | PCA1N-A | 563450 | 5.69 | 0.17 | 0.61 | 190 | 0.018 | 5.794 | 0.148 | 2.552 | Study
C6000-B:C2 | GLYCATED HAEMOGLOBIN (NGSP) | PCA1N | 617428 | 5.7 | 0.34 | 0.178 | 357 | 0.011 | 5.761 | 0.109 | 1.891 | In use
C6000-B:C2 | GLYCATED HAEMOGLOBIN (NGSP) | PCA1N | 548925 | 5.65 | 0.17 | 0.478 | 422 | 0.014 | 5.731 | 0.094 | 1.64 | Study
C6000-B:C2 | GLYCATED HAEMOGLOBIN (NGSP) | PCA1P | 509113 | 10.7 | 0.32 | 0.5 | 110 | 0.015 | 10.86 | 0.122 | 1.126 | Study
C6000-A:C1 | GLYCATED HAEMOGLOBIN (NGSP) | PCA1P-A | 563451 | 10.6 | 0.32 | -0.041 | 191 | -0.001 | 10.587 | 0.253 | 2.393 | Study
C6000-B:C2 | HAPTOGLOBIN | s-PCCC 2-B | 461597 | 1.37 | 0.11 | -0.563 | 36 | -0.045 | 1.308 | 0.018 | 1.381 | Study
C6000-B:C2 | HAPTOGLOBIN | s-PCCC 1-B | 525027 | 0.804 | 0.064 | -0.613 | 43 | -0.049 | 0.765 | 0.009 | 1.216 | Study
C6000-B:C2 | HAPTOGLOBIN | PCCC 2-B | 519198 | 1.24 | 0.1 | -0.191 | 332 | -0.015 | 1.221 | 0.036 | 2.94 | Study
C6000-B:C2 | HAPTOGLOBIN | PCCC 1-B | 525205 | 0.765 | 0.061 | 0.569 | 156 | 0.045 | 0.8 | 0.018 | 2.256 | In use
C6000-B:C2 | HAPTOGLOBIN | PCCC 2-B | 535719 | 1.21 | 0.1 | 0.435 | 104 | 0.036 | 1.253 | 0.016 | 1.312 | In use
C6000-B:C2 | HAPTOGLOBIN | PCCC 1-B | 525027 | 0.804 | 0.064 | -0.349 | 284 | -0.028 | 0.782 | 0.022 | 2.87 | Study
C6000-A:C1 | HDL CHOLESTEROL | PCCC 1 | 525027 | 0.762 | 0.061 | 0.038 | 283 | 0.003 | 0.764 | 0.024 | 3.124 | Study
C6000-A:C1 | HDL CHOLESTEROL | PCCC 2 | 519198 | 1.46 | 0.12 | -0.027 | 316 | -0.002 | 1.457 | 0.043 | 2.962 | Study
C6000-A:C1 | HDL CHOLESTEROL | PCCC 2 | 535719 | 1.52 | 0.12 | 0.626 | 84 | 0.049 | 1.595 | 0.064 | 4.044 | In use
C6000-A:C1 | HDL CHOLESTEROL | s-PCCC 2 | 461597 | 1.41 | 0.11 | 0.188 | 32 | 0.015 | 1.431 | 0.022 | 1.527 | Study
C6000-A:C1 | HDL CHOLESTEROL | PCCC 1 | 525205 | 0.719 | 0.058 | 0.581 | 131 | 0.047 | 0.753 | 0.037 | 4.857 | In use
C6000-A:C1 | HDL CHOLESTEROL | s-PCCC 1 | 525027 | 0.762 | 0.061 | 0.367 | 39 | 0.029 | 0.784 | 0.017 | 2.145 | Study
C6000-B:E3-1 | HIGH-SENSITIVITY TROPONIN T | PC TN1-B | 590216 | 26.3 | 2.37 | 0.634 | 157 | 0.057 | 27.803 | 1.637 | 5.887 | Study
C6000-A:E2-1 | HIGH-SENSITIVITY TROPONIN T | Ran Trop T | 4607CK | 6.38 | 0.64 | 0.147 | 198 | 0.015 | 6.474 | 0.688 | 10.625 | In use
C6000-B:E3-1 | HIGH-SENSITIVITY TROPONIN T | PC TN1-B | 558106 | 25.1 | 2.26 | -0.378 | 86 | -0.034 | 24.246 | 0.496 | 2.047 | Study
C6000-A:E2-1 | HIGH-SENSITIVITY TROPONIN T | PC TN1 | 590216 | 26.3 | 2.37 | 0.595 | 163 | 0.054 | 27.711 | 0.85 | 3.068 | Study
C6000-A:E2-1 | HIGH-SENSITIVITY TROPONIN T | BIO CTC1 | 89041 | 17.99 | 0.56 | 0.351 | 160 | 0.011 | 18.187 | 0.714 | 3.927 | In use
C6000-A:E2-1 | HIGH-SENSITIVITY TROPONIN T | PC TN1 | 558106 | 25.1 | 2.26 | 0.216 | 82 | 0.019 | 25.588 | 0.697 | 2.724 | Study
C6000-B:E3-1 | HIGH-SENSITIVITY TROPONIN T | BIO CTC1-B | 89041 | 17.99 | 0.56 | 1.683 | 181 | 0.052 | 18.932 | 0.816 | 4.309 | In use
C6000-B:E3-1 | HIGH-SENSITIVITY TROPONIN T | Ran Trop T-B | 4607CK | 6.38 | 0.65 | -0.146 | 197 | -0.015 | 6.285 | 0.616 | 9.793 | Study
C6000-A:E2-1 | HIGH-SENSITIVITY TROPONIN T | PC TN1 | 629183 | 28.6 | 2.57 | 0.055 | 164 | 0.005 | 28.741 | 0.95 | 3.304 | Study
C6000-B:E3-1 | HIGH-SENSITIVITY TROPONIN T | PC TN1-B | 629183 | 28.6 | 2.57 | 0.526 | 176 | 0.047 | 29.951 | 1.212 | 4.046 | Study
C6000-A:E2-1 | HIGH-SENSITIVITY TROPONIN T | BIO CTC1 | 89021 | 17.72 | 1.267 | 0.368 | 53 | 0.026 | 18.186 | 0.482 | 2.653 | Study
C6000-B:E3-1 | HIGH-SENSITIVITY TROPONIN T | BIO CTC1-B | 89021 | 17.72 | 1.267 | -0.432 | 49 | -0.031 | 17.173 | 0.36 | 2.098 | Study
LIAISON XL | Human growth hormone | hGH L2 | 9260310 | 17.8 | 1.2 | -0.292 | 52 | -0.02 | 17.45 | 0.623 | 3.568 | Study
LIAISON XL | Human growth hormone | hGH L1 | 7980310 | 3.81 | 0.257 | -0.154 | 55 | -0.01 | 3.771 | 0.148 | 3.914 | Study
C6000-A:C1 | IGG CSF | PNPUC TPC3 | 649811 | 21.5 | 1.7 | -0.129 | 157 | -0.01 | 21.281 | 0.911 | 4.28 | Study
C6000-A:C1 | IGG CSF | PNPUC TPC3 | 531179 | 20 | 1.6 | 0.019 | 75 | 0.002 | 20.031 | 1.022 | 5.101 | Study
C6000-A:C1 | IGG CSF | PPPUC TPC3 | 531181 | 140 | 11 | -0.531 | 32 | -0.042 | 134.157 | 4.527 | 3.374 | Study
C6000-A:C1 | IGG CSF | PNPUC TPC3 | 602712 | 23.7 | 1.9 | -0.367 | 109 | -0.029 | 23.003 | 1.447 | 6.289 | Study
C6000-A:C1 | IGG CSF | PNPUC TPC3 | 673292 | 21.2 | 1.7 | -0.717 | 35 | -0.057 | 19.981 | 0.95 | 4.753 | In use
C6000-A:C1 | IGG CSF | PPPUC TPC3 | 602710 | 135 | 11 | -0.136 | 154 | -0.011 | 133.509 | 4.377 | 3.278 | Study
C6000-A:C1 | IGG CSF | PPPUC TPC3 | 637091 | 142 | 11 | 0.033 | 101 | 0.003 | 142.362 | 3.608 | 2.534 | Study
C6000-A:C1 | IGG CSF | PPPUC TPC3 | 646927 | 129 | 10 | -0.09 | 112 | -0.007 | 128.098 | 4.895 | 3.821 | In use
C6000-B:C2 | IMMUNOGLOBULIN A | PCCC 2-B | 535719 | 2.26 | 0.18 | -0.145 | 63 | -0.012 | 2.234 | 0.028 | 1.252 | In use
C6000-B:C2 | IMMUNOGLOBULIN A | PCCC 1-B | 525027 | 1.51 | 0.12 | 0.229 | 195 | 0.018 | 1.538 | 0.045 | 2.907 | Study
C6000-B:C2 | IMMUNOGLOBULIN A | PCCC 1-B | 525205 | 1.54 | 0.12 | -0.352 | 103 | -0.027 | 1.498 | 0.037 | 2.443 | In use
C6000-B:C2 | IMMUNOGLOBULIN A | PCCC 2-B | 519198 | 2.42 | 0.19 | 0.202 | 228 | 0.016 | 2.458 | 0.06 | 2.422 | Study
C6000-B:C2 | IMMUNOGLOBULIN G | PCCC 2-B | 519198 | 11.8 | 0.7 | 0.201 | 226 | 0.012 | 11.94 | 0.318 | 2.662 | Study
C6000-B:C2 | IMMUNOGLOBULIN G | PCCC 1-B | 525205 | 7.35 | 0.44 | 0.767 | 108 | 0.046 | 7.688 | 0.193 | 2.515 | In use
C6000-B:C2 | IMMUNOGLOBULIN G | PCCC 2-B | 535719 | 11.2 | 0.7 | 0.642 | 68 | 0.04 | 11.649 | 0.274 | 2.352 | In use
C6000-B:C2 | IMMUNOGLOBULIN G | PCCC 1-B | 525027 | 7.27 | 0.44 | 0.255 | 193 | 0.015 | 7.382 | 0.212 | 2.878 | Study
C6000-B:C2 | IMMUNOGLOBULIN G | s-PCCC 1-B | 525027 | 7.27 | 0.44 | 0.38 | 32 | 0.023 | 7.437 | 0.256 | 3.439 | Study
C6000-B:C2 | IMMUNOGLOBULIN M | PCCC 1-B | 525205 | 0.623 | 0.05 | 1.006 | 109 | 0.081 | 0.673 | 0.027 | 4.026 | In use
C6000-B:C2 | IMMUNOGLOBULIN M | PCCC 2-B | 535719 | 1.02 | 0.08 | 0.483 | 67 | 0.038 | 1.059 | 0.024 | 2.29 | In use
C6000-B:C2 | IMMUNOGLOBULIN M | PCCC 2-B | 519198 | 1.04 | 0.08 | 0.628 | 236 | 0.048 | 1.09 | 0.04 | 3.663 | Study
C6000-B:C2 | IMMUNOGLOBULIN M | PCCC 1-B | 525027 | 0.625 | 0.05 | 0.76 | 199 | 0.061 | 0.663 | 0.033 | 4.925 | Study
C6000-B:C2 | INORGANIC PHOSPHATE URINE | Ran-Liq 2 | 1209UC | 8.4 | 0.84 | 0.398 | 213 | 0.04 | 8.734 | 0.213 | 2.433 | Study
C6000-B:C2 | INORGANIC PHOSPHATE URINE | Bio-liq1 | 88161 | 8.16 | 0.435 | 0.326 | 113 | 0.017 | 8.302 | 0.17 | 2.051 | Study
C6000-B:C2 | INORGANIC PHOSPHATE URINE | Bio-liq 2 | 88162 | 16.2 | 0.75 | 0.3 | 107 | 0.014 | 16.425 | 0.388 | 2.364 | Study
C6000-B:C2 | INORGANIC PHOSPHATE URINE | Ran-Liq 3 | 1238UC | 13.8 | 1.4 | 0.593 | 127 | 0.06 | 14.63 | 0.336 | 2.297 | In use
C6000-B:C2 | INORGANIC PHOSPHATE URINE | Ran-Liq 3 | 1210UC | 13.9 | 1.4 | 0.346 | 214 | 0.035 | 14.385 | 0.321 | 2.232 | Study
C6000-B:C2 | INORGANIC PHOSPHATE URINE | Ran-Liq 2 | 1237UC | 8.23 | 0.83 | 0.593 | 115 | 0.06 | 8.722 | 0.179 | 2.052 | In use
C6000-A:E1-2 | INSULIN | RIAP L3 | 2110EC | 27.9 | 3.5 | -0.023 | 156 | -0.003 | 27.819 | 1.596 | 5.737 | Study
C6000-A:E1-2 | INSULIN | RIAP L3 | 2147 | 26.4 | 3.3 | 0.06 | 164 | 0.007 | 26.598 | 0.667 | 2.509 | Study
C6000-A:E1-2 | INSULIN | RIAP L3 | 2181EC | 28.8 | 3.6 | -0.163 | 165 | -0.02 | 28.213 | 1.274 | 4.514 | In use
C6000-A:E1-2 | INSULIN | RIAP L2 | 2228EC | 16.1 | 2 | -0.067 | 81 | -0.008 | 15.966 | 1.129 | 7.071 | In use
C6000-A:E1-2 | INSULIN | RIAP L2 | 2124EC | 12 | 1.5 | -0.097 | 98 | -0.012 | 11.855 | 1.263 | 10.657 | Study
C6000-A:E1-2 | INSULIN | RIAP L2 | 2107EC | 13.9 | 1.75 | 0.015 | 316 | 0.002 | 13.926 | 0.682 | 4.897 | Study
LIAISON XL | Insulin-like growth factor I | IGF-1 L2 | 1422081 | 253 | 32.667 | -0.603 | 36 | -0.078 | 233.297 | 41.319 | 17.711 | Study
LIAISON XL | Insulin-like growth factor I | IGF-1 L2 | 1422084 | 253 | 29.333 | -0.504 | 37 | -0.058 | 238.219 | 37.909 | 15.914 | In use
LIAISON XL | Insulin-like growth factor I | IGF-1 L1 | 1412084 | 54 | 7.233 | 0.861 | 35 | 0.115 | 60.231 | 8.127 | 13.493 | In use
C6000-A:C1 | IRON | PCCC 2 | 461597 | 43.1 | 2.6 | 0.266 | 67 | 0.016 | 43.791 | 0.712 | 1.625 | Study
C6000-A:C1 | IRON | PCCC 2 | 535719 | 44.2 | 2.7 | -0.189 | 93 | -0.012 | 43.69 | 0.784 | 1.794 | In use
C6000-A:C1 | IRON | PCCC 2 | 519198 | 43 | 2.6 | 0.06 | 375 | 0.004 | 43.156 | 0.877 | 2.031 | Study
C6000-A:C1 | IRON | PCCC 1 | 525205 | 19.3 | 1.2 | 0.174 | 152 | 0.011 | 19.509 | 0.469 | 2.404 | In use
C6000-A:C1 | IRON | PCCC 1 | 525027 | 19.6 | 1.2 | 0.243 | 401 | 0.015 | 19.892 | 0.654 | 3.288 | Study
C6000-B:C2 | LACTATE | PCCC 1-B | 525027 | 1.8 | 0.11 | -0.359 | 231 | -0.022 | 1.76 | 0.052 | 2.952 | Study
C6000-B:C2 | LACTATE | PCCC 2-B | 535719 | 3.82 | 0.23 | -0.038 | 83 | -0.002 | 3.811 | 0.056 | 1.474 | In use
C6000-B:C2 | LACTATE | s-PCCC 1-B | 525027 | 1.8 | 0.11 | -0.419 | 33 | -0.026 | 1.754 | 0.027 | 1.548 | Study
C6000-B:C2 | LACTATE | PCCC 1-B | 525205 | 1.78 | 0.11 | 0.006 | 135 | 0 | 1.781 | 0.056 | 3.131 | In use
C6000-B:C2 | LACTATE | s-PCCC 2-B | 461597 | 3.79 | 0.23 | -0.572 | 30 | -0.035 | 3.658 | 0.034 | 0.942 | Study
C6000-B:C2 | LACTATE | PCCC 2-B | 519198 | 3.81 | 0.23 | 0.035 | 271 | 0.002 | 3.818 | 0.105 | 2.754 | Study
C6000-A:C1 | LACTATE DEHYDROGENASE | s-PCCC 1 | 525027 | 175 | 11 | 0.01 | 47 | 0.001 | 175.106 | 1.992 | 1.137 | Study
C6000-A:C1 | LACTATE DEHYDROGENASE | PCCC 2 | 519198 | 310 | 19 | -0.071 | 361 | -0.004 | 308.645 | 6.764 | 2.192 | Study
C6000-B:C2 | LACTATE DEHYDROGENASE | PCCC 2-B | 535719 | 299 | 18 | -0.086 | 88 | -0.005 | 297.455 | 7.523 | 2.529 | In use
C6000-B:C2 | LACTATE DEHYDROGENASE | PCCC 2-B | 519198 | 310 | 19 | 0.144 | 295 | 0.009 | 312.736 | 9.699 | 3.101 | Study
C6000-B:C2 | LACTATE DEHYDROGENASE | PCCC 1-B | 525027 | 175 | 11 | 0.031 | 256 | 0.002 | 175.336 | 4.48 | 2.555 | Study
C6000-B:C2 | LACTATE DEHYDROGENASE | PCCC 1-B | 525205 | 164 | 10 | 0.216 | 138 | 0.013 | 166.159 | 6.291 | 3.786 | In use
C6000-A:C1 | LACTATE DEHYDROGENASE | PCCC 2 | 535719 | 299 | 18 | 0.133 | 98 | 0.008 | 301.398 | 6.916 | 2.295 | In use
C6000-B:C2 | LACTATE DEHYDROGENASE | s-PCCC 1-B | 525027 | 175 | 11 | -0.021 | 39 | -0.001 | 174.769 | 2.67 | 1.528 | Study
C6000-A:C1 | LACTATE DEHYDROGENASE | s-PCCC 2 | 461597 | 311 | 12.9 | -0.048 | 40 | -0.002 | 310.375 | 3.216 | 1.036 | Study
C6000-A:C1 | LACTATE DEHYDROGENASE | PCCC 1 | 525027 | 175 | 11 | -0.131 | 321 | -0.008 | 173.555 | 5.216 | 3.006 | Study
C6000-A:C1 | LACTATE DEHYDROGENASE | PCCC 1 | 525205 | 164 | 10 | 0.132 | 155 | 0.008 | 165.316 | 3.862 | 2.336 | In use
C6000-B:C2 | LACTATE DEHYDROGENASE | s-PCCC 2-B | 461597 | 311 | 12.9 | -0.164 | 34 | -0.007 | 308.882 | 4.395 | 1.423 | Study
C6000-A:C1 | LIPASE | s-PCCC 2 | 461597 | 95.4 | 5.7 | -0.29 | 43 | -0.017 | 93.747 | 2.927 | 3.123 | Study
C6000-A:C1 | LIPASE | PCCC 2 | 535719 | 100 | 6 | -0.42 | 100 | -0.025 | 97.478 | 3.05 | 3.128 | In use
C6000-A:C1 | LIPASE | PCCC 1 | 525027 | 45.3 | 2.7 | -0.342 | 338 | -0.02 | 44.378 | 1.574 | 3.547 | Study
C6000-A:C1 | LIPASE | s-PCCC 1 | 525027 | 45.3 | 2.7 | -0.585 | 47 | -0.035 | 43.721 | 1.009 | 2.308 | Study
C6000-A:C1 | LIPASE | PCCC 2 | 519198 | 101 | 6 | -0.069 | 382 | -0.004 | 100.585 | 3.579 | 3.558 | Study
C6000-A:C1 | LIPASE | PCCC 1 | 525205 | 45 | 2.7 | -0.346 | 161 | -0.021 | 44.065 | 1.326 | 3.01 | In use
C6000-B:C2 | LIPOPROTEIN A2 | PC LPA 1 | 520719 | 44.8 | 3 | -0.78 | 95 | -0.052 | 42.46 | 2.042 | 4.809 | In use
C6000-B:C2 | LIPOPROTEIN A2 | PC LPA 2 | 520720 | 112 | 7 | -0.464 | 92 | -0.029 | 108.754 | 6.695 | 6.156 | In use
C6000-B:C2 | LITHIUM | s-PCCC 1-B | 525027 | 0.912 | 0.036 | 0.755 | 37 | 0.03 | 0.939 | 0.025 | 2.613 | Study
C6000-B:C2 | LITHIUM | PCCC 2-B | 519198 | 1.8 | 0.07 | -0.092 | 276 | -0.004 | 1.794 | 0.056 | 3.096 | Study
C6000-B:C2 | LITHIUM | PCCC 1-B | 525205 | 0.935 | 0.037 | -0.427 | 135 | -0.017 | 0.919 | 0.04 | 4.362 | In use
C6000-B:C2 | LITHIUM | s-PCCC 2-B | 461597 | 1.79 | 0.07 | 0.084 | 34 | 0.003 | 1.796 | 0.046 | 2.554 | Study
C6000-B:C2 | LITHIUM | PCCC 2-B | 535719 | 1.81 | 0.07 | -0.212 | 85 | -0.008 | 1.795 | 0.067 | 3.705 | In use
C6000-B:C2 | LITHIUM | PCCC 1-B | 525027 | 0.935 | 0.037 | -0.717 | 243 | -0.028 | 0.908 | 0.04 | 4.37 | Study
C6000-A:E1-2 | LUTEINISING HORMONE | RIAP L3 | 2181EC | 55.02 | 1.96 | -0.215 | 132 | -0.008 | 54.598 | 2.554 | 4.678 | In use
C6000-A:E1-2 | LUTEINISING HORMONE | RIAP L3 | 2110EC | 55.6 | 5.55 | -0.383 | 136 | -0.038 | 53.477 | 2.135 | 3.993 | Study
C6000-A:E1-2 | LUTEINISING HORMONE | RIAP L1 | 2105EC | 3.669 | 0.221 | -0.42 | 298 | -0.025 | 3.576 | 0.127 | 3.565 | Study
C6000-A:E1-2 | LUTEINISING HORMONE | RIAP L1 | 2179EC | 4.26 | 0.327 | -1.565 | 125 | -0.12 | 3.748 | 0.219 | 5.852 | In use
C6000-A:E1-2 | LUTEINISING HORMONE | RIAP L3 | 2147 | 54.7 | 5.45 | 0.216 | 154 | 0.022 | 55.878 | 1.918 | 3.432 | Study
C6000-B:C2 | MAGNESIUM | PCCC 2-B | 535719 | 1.36 | 0.05 | -0.289 | 99 | -0.011 | 1.346 | 0.05 | 3.716 | In use
C6000-A:C1 | MAGNESIUM | PCCC 2 | 461597 | 1.28 | 0.05 | 0.067 | 71 | 0.003 | 1.283 | 0.017 | 1.348 | Study
C6000-B:C2 | MAGNESIUM | PCCC 1-B | 525205 | 0.793 | 0.032 | -0.083 | 168 | -0.003 | 0.79 | 0.031 | 3.907 | In use
C6000-B:C2 | MAGNESIUM | PCCC 2-B | 461597 | 1.28 | 0.05 | 0.279 | 71 | 0.011 | 1.294 | 0.022 | 1.677 | Study
C6000-A:C1 | MAGNESIUM | PCCC 1 | 525027 | 0.791 | 0.032 | 0.019 | 417 | 0.001 | 0.792 | 0.019 | 2.373 | Study
C6000-B:C2 | MAGNESIUM | PCCC 1-B | 525027 | 0.791 | 0.043 | 0.06 | 388 | 0.003 | 0.794 | 0.022 | 2.725 | Study
C6000-A:C1 | MAGNESIUM | PCCC 1 | 525205 | 0.793 | 0.032 | -0.106 | 165 | -0.004 | 0.79 | 0.018 | 2.317 | In use
C6000-A:C1 | MAGNESIUM | PCCC 2 | 535719 | 1.36 | 0.05 | -0.23 | 103 | -0.008 | 1.348 | 0.029 | 2.167 | In use
C6000-B:C2 | MAGNESIUM | PCCC 2-B | 519198 | 1.35 | 0.05 | 0.107 | 370 | 0.004 | 1.355 | 0.032 | 2.351 | Study
C6000-A:C1 | MAGNESIUM | PCCC 2 | 519198 | 1.35 | 0.05 | 0 | 389 | 0 | 1.35 | 0.028 | 2.091 | Study
C6000-B:C2 | MAGNESIUM URINE | Ran-Liq 2 | 1209UC | 2.46 | 0.25 | 0.681 | 215 | 0.069 | 2.63 | 0.312 | 11.858 | Study
C6000-B:C2 | MAGNESIUM URINE | Ran-Liq 3 | 1210UC | 5.13 | 0.52 | 0.422 | 216 | 0.043 | 5.35 | 0.34 | 6.355 | Study
C6000-B:C2 | MAGNESIUM URINE | Ran-Liq 3 | 1238UC | 6.17 | 0.62 | 0.396 | 134 | 0.04 | 6.415 | 0.214 | 3.342 | In use
C6000-B:C2 | MAGNESIUM URINE | Bio-liq 2 | 88162 | 3.83 | 0.202 | 0.469 | 125 | 0.025 | 3.925 | 0.114 | 2.894 | Study
C6000-B:C2 | MAGNESIUM URINE | Ran-Liq 2 | 1237UC | 2.88 | 0.29 | 0.728 | 122 | 0.073 | 3.091 | 0.119 | 3.845 | In use
C6000-B:C2 | MAGNESIUM URINE | Bio-liq1 | 88161 | 1.64 | 0.09 | 0.132 | 126 | 0.007 | 1.652 | 0.105 | 6.358 | Study
C6000-A:E1-1 | OESTRADIOL III | RIAP L1 | 2179EC | 198 | 20 | 0.192 | 130 | 0.019 | 201.844 | 15.591 | 7.724 | In use
C6000-A:E1-1 | OESTRADIOL III | RIAP L3 | 2110EC | 1241 | 124 | 0.039 | 143 | 0.004 | 1245.881 | 47.338 | 3.8 | Study
C6000-A:E1-1 | OESTRADIOL III | RIAP L3 | 2147 | 1460 | 146 | 0.323 | 157 | 0.032 | 1507.191 | 47.681 | 3.164 | Study
C6000-A:E1-1 | OESTRADIOL III | RIAP L3 | 2181EC | 1480 | 148 | 0.219 | 133 | 0.022 | 1512.406 | 35.75 | 2.364 | In use
C6000-A:E1-1 | OESTRADIOL III | RIAP L1 | 2105EC | 169 | 24.4 | 0.313 | 303 | 0.045 | 176.634 | 19.59 | 11.091 | Study
OSMO1 | Osmolality | Serim BIOLIQ 1 | 45871 | 306 | 5.5 | -0.003 | 240 | 0 | 305.983 | 2.778 | 0.908 | In use
OSMO1 | Osmolality | Serim BIOLIQ 1 | 45931 | 311 | 3.833 | 0.206 | 114 | 0.003 | 311.789 | 2.798 | 0.898 | In use
OSMO1 | Osmolality | Serim BIOLIQ 2 | 45952 | 398 | 9.3 | 0.24 | 74 | 0.006 | 400.23 | 4.586 | 1.146 | In use
OSMO1 | Osmolality | Serim BIOLIQ 2 | 45872 | 410 | 4 | -0.407 | 110 | -0.004 | 408.373 | 3.199 | 0.783 | In use
OSMO1 | Osmolality | Serim BIOLIQ 2 | 45912 | 416 | 5.833 | -0.776 | 139 | -0.011 | 411.475 | 3.552 | 0.863 | Study
OSMO1 | Osmolality | Serim BIOLIQ 2 | 45862 | 401 | 4 | 0.061 | 41 | 0.001 | 401.244 | 3.145 | 0.784 | Study
C6000-B:C2 | PARACETAMOL 2 | RIAP L2-B | 2107EC | 320 | 32 | 0 | 280 | 0 | 319.991 | 15.205 | 4.752 | Study
C6000-B:C2 | PARACETAMOL 2 | RIAP L2-B | 2228EC | 323 | 21.7 | 0.05 | 70 | 0.003 | 324.084 | 20.702 | 6.388 | In use
C6000-B:C2 | PARACETAMOL 2 | RIAP L1-B | 2105EC | 152 | 7.6 | 0.049 | 288 | 0.002 | 152.374 | 6.371 | 4.181 | Study
C6000-B:C2 | PARACETAMOL 2 | RIAP L2-B | 2124EC | 312 | 15.5 | 0.679 | 77 | 0.034 | 322.526 | 14.722 | 4.565 | Study
C6000-B:C2 | PARACETAMOL 2 | RIAP L1-B | 2179EC | 153 | 15.5 | 0.07 | 123 | 0.007 | 154.079 | 8.861 | 5.751 | In use
C6000-A:E2-1 | PARATHYROID HORMONE | PC V2 | 581988 | 19.6 | 1.57 | 0.802 | 55 | 0.064 | 20.859 | 0.286 | 1.373 | Study
C6000-A:E2-1 | PARATHYROID HORMONE | PC V1 | 612148 | 6.23 | 0.56 | 0.244 | 66 | 0.022 | 6.367 | 0.197 | 3.088 | In use
C6000-A:E2-1 | PARATHYROID HORMONE | PC V1 | 581987 | 6 | 0.54 | 0.864 | 55 | 0.078 | 6.467 | 0.268 | 4.138 | Study
C6000-A:E2-1 | PARATHYROID HORMONE | PC V2 | 573356 | 21.2 | 1.7 | -0.696 | 92 | -0.056 | 20.016 | 0.816 | 4.076 | Study
C6000-A:E2-1 | PARATHYROID HORMONE | PC V1 | 573355 | 6.11 | 0.55 | 0.052 | 92 | 0.005 | 6.138 | 0.276 | 4.493 | Study
C6000-A:E2-1 | PARATHYROID HORMONE | PC V1 | 595763 | 6.6 | 0.59 | 0.116 | 209 | 0.01 | 6.668 | 0.196 | 2.937 | Study
C6000-A:E2-1 | PARATHYROID HORMONE | PC V2 | 595765 | 19.1 | 1.53 | 0.579 | 235 | 0.046 | 19.985 | 0.809 | 4.048 | In use
C6000-A:C1 | PHOSPHATE | PCCC 2 | 461597 | 2.51 | 0.13 | 0.084 | 72 | 0.004 | 2.521 | 0.039 | 1.531 | Study
C6000-A:C1 | PHOSPHATE | PCCC 1 | 525027 | 1.3 | 0.07 | -0.15 | 417 | -0.008 | 1.29 | 0.038 | 2.974 | Study
C6000-B:C2 | PHOSPHATE | PCCC 2-B | 535719 | 2.76 | 0.14 | -0.276 | 89 | -0.014 | 2.721 | 0.054 | 1.968 | In use
C6000-A:C1 | PHOSPHATE | PCCC 2 | 519198 | 2.58 | 0.13 | -0.279 | 389 | -0.014 | 2.544 | 0.049 | 1.909 | Study
C6000-B:C2 | PHOSPHATE | PCCC 1-B | 525027 | 1.3 | 0.07 | -0.178 | 366 | -0.01 | 1.288 | 0.026 | 1.988 | Study
C6000-B:C2 | PHOSPHATE | PCCC 2-B | 519198 | 2.58 | 0.13 | -0.123 | 346 | -0.006 | 2.564 | 0.058 | 2.281 | Study
C6000-B:C2 | PHOSPHATE | PCCC 1-B | 525205 | 1.19 | 0.06 | 0.001 | 151 | 0 | 1.19 | 0.038 | 3.172 | In use
C6000-A:C1 | PHOSPHATE | PCCC 1 | 525205 | 1.19 | 0.06 | -0.511 | 164 | -0.026 | 1.159 | 0.024 | 2.05 | In use
C6000-A:C1 | PHOSPHATE | PCCC 2 | 535719 | 2.76 | 0.14 | -0.495 | 104 | -0.025 | 2.691 | 0.034 | 1.271 | In use
C6000-B:C2 | PHOSPHATE | PCCC 2-B | 461597 | 2.51 | 0.13 | -0.354 | 70 | -0.018 | 2.464 | 0.03 | 1.202 | Study
C6000-A:ISE1 | POTASSIUM | s-PCCC 2 | 535719 | 7.26 | 0.22 | 0.605 | 192 | 0.018 | 7.393 | 0.121 | 1.631 | In use
C6000-B:ISE1 | POTASSIUM | s-PCCC 2-B | 535719 | 7.26 | 0.22 | 0.207 | 190 | 0.006 | 7.306 | 0.097 | 1.328 | In use
C6000-B:ISE1 | POTASSIUM | PCCC 2-B | 461597 | 6.87 | 0.21 | -0.137 | 92 | -0.004 | 6.841 | 0.055 | 0.81 | Study
C6000-B:ISE1 | POTASSIUM | s-PCCC 1-B | 525027 | 3.64 | 0.11 | -0.012 | 565 | 0 | 3.639 | 0.067 | 1.833 | Study
C6000-A:ISE1 | POTASSIUM | PCCC 1 | 525027 | 3.64 | 0.11 | 0.087 | 106 | 0.003 | 3.65 | 0.06 | 1.654 | Study
C6000-B:ISE1 | POTASSIUM | PCCC 1-B | 525027 | 3.64 | 0.11 | 0.125 | 368 | 0.004 | 3.654 | 0.059 | 1.601 | Study
C6000-A:ISE1 | POTASSIUM | PCCC 2 | 461597 | 6.87 | 0.21 | -0.176 | 88 | -0.005 | 6.833 | 0.081 | 1.19 | Study
C6000-A:ISE1 | POTASSIUM | s-PCCC 1 | 525205 | 3.49 | 0.1 | 0.964 | 330 | 0.028 | 3.586 | 0.08 | 2.236 | In use
C6000-A:ISE1 | POTASSIUM | s-PCCC 2 | 519198 | 7.06 | 0.21 | 0.424 | 717 | 0.013 | 7.149 | 0.119 | 1.669 | Study
C6000-B:ISE1 | POTASSIUM | PCCC 2-B | 519198 | 7.06 | 0.21 | 0.311 | 280 | 0.009 | 7.125 | 0.092 | 1.289 | Study
C6000-B:ISE1 | POTASSIUM | s-PCCC 2-B | 519198 | 7.06 | 0.21 | 0.252 | 686 | 0.007 | 7.113 | 0.111 | 1.56 | Study
C6000-A:ISE1 | POTASSIUM | s-PCCC 1 | 525027 | 3.64 | 0.11 | 0.043 | 597 | 0.001 | 3.645 | 0.07 | 1.915 | Study
C6000-B:ISE1 | POTASSIUM | s-PCCC 1-B | 525205 | 3.49 | 0.1 | 0.566 | 293 | 0.016 | 3.547 | 0.095 | 2.665 | In use
C6000-B:ISE1 | POTASSIUM URINE | Ran-Liq 2 | 1209UC | 29.2 | 2.2 | -0.044 | 247 | -0.003 | 29.103 | 0.955 | 3.281 | Study
C6000-B:ISE1 | POTASSIUM URINE | Ran-Liq 3 | 1210UC | 71 | 5.3 | 0.1 | 249 | 0.007 | 71.532 | 3.907 | 5.462 | Study
C6000-B:ISE1 | POTASSIUM URINE | Bio-liq 2 | 88162 | 66.1 | 2.45 | 0.546 | 137 | 0.02 | 67.437 | 1.865 | 2.765 | Study
C6000-B:ISE1 | POTASSIUM URINE | Bio-liq1 | 88161 | 30.2 | 0.9 | 0.692 | 144 | 0.021 | 30.823 | 0.606 | 1.968 | Study
C6000-B:ISE1 | POTASSIUM URINE | Ran-Liq 2 | 1237UC | 27.5 | 2.05 | 0.771 | 137 | 0.057 | 29.08 | 0.792 | 2.725 | In use
C6000-B:ISE1 | POTASSIUM URINE | Ran-Liq 3 | 1238UC | 72.38 | 1.48 | 0.353 | 149 | 0.007 | 72.903 | 2.668 | 3.66 | In use
C6000-A:E1-1 | proBNP | PC CARD1 | 657390 | 127 | 7.62 | -0.108 | 65 | -0.006 | 126.177 | 5.442 | 4.313 | Study
C6000-A:E1-1 | proBNP | PC CARD2 | 657389 | 4620 | 277 | -0.491 | 63 | -0.029 | 4484.127 | 183.837 | 4.1 | Study
C6000-A:E1-1 | proBNP | PC CARD1 | 629283 | 132 | 7.92 | -0.465 | 63 | -0.028 | 128.321 | 4.909 | 3.826 | In use
C6000-A:E1-1 | proBNP | PC CARD1 | 586866 | 125 | 7.5 | -0.528 | 87 | -0.032 | 121.037 | 4.757 | 3.93 | Study
C6000-A:E1-1 | proBNP | PC CARD2 | 629284 | 4930 | 296 | -0.63 | 61 | -0.038 | 4743.607 | 156.691 | 3.303 | In use
C6000-A:E1-1 | proBNP | PC CARD1 | 623415 | 127 | 6.07 | -0.575 | 163 | -0.028 | 123.507 | 4.025 | 3.259 | Study
C6000-A:E1-1 | proBNP | PC CARD2 | 623416 | 4868 | 211 | -0.665 | 162 | -0.029 | 4727.735 | 141.411 | 2.991 | Study
C6000-A:E1-1 | proBNP | PC CARD2 | 586867 | 4870 | 292 | 0.005 | 87 | 0 | 4871.494 | 187.237 | 3.844 | Study
C6000-A:E1-2 | PROCALCITONIN (SENSITIVE) | PC PCT1 | 577640 | 0.48 | 0.03 | -0.607 | 33 | -0.038 | 0.462 | 0.015 | 3.161 | Study
C6000-A:E1-2 | PROCALCITONIN (SENSITIVE) | PC PCT2 | 577646 | 9.18 | 0.55 | -0.677 | 33 | -0.041 | 8.808 | 0.355 | 4.033 | Study
C6000-A:E2-1 | PROCALCITONIN (SENSITIVE) | PC PCT2 | 671599 | 9.39 | 0.56 | -0.553 | 107 | -0.033 | 9.08 | 0.192 | 2.113 | In use
C6000-A:E2-1 | PROCALCITONIN (SENSITIVE) | PC PCT2 | 614523 | 9.98 | 0.6 | -0.684 | 61 | -0.041 | 9.569 | 0.255 | 2.66 | Study
C6000-A:E2-1 | PROCALCITONIN (SENSITIVE) | PC PCT1 | 671597 | 0.49 | 0.03 | 0.047 | 108 | 0.003 | 0.491 | 0.014 | 2.904 | In use
C6000-A:E2-1 | PROCALCITONIN (SENSITIVE) | PC PCT1 | 614519 | 0.51 | 0.04 | -0.341 | 61 | -0.027 | 0.496 | 0.012 | 2.425 | Study
C6000-A:E2-1 | PROCALCITONIN (SENSITIVE) | PC PCT2 | 559763 | 9.5 | 0.57 | -0.142 | 182 | -0.009 | 9.419 | 0.379 | 4.027 | Study
C6000-A:E2-1 | PROCALCITONIN (SENSITIVE) | PC PCT1 | 559760 | 0.5 | 0.04 | 0.227 | 181 | 0.018 | 0.509 | 0.024 | 4.799 | Study
C6000-A:E2-2 | PROGESTERONE III | RIAP L1 | 2105EC | 3.47 | 0.43 | -0.213 | 287 | -0.026 | 3.378 | 0.233 | 6.885 | Study
C6000-A:E2-2 | PROGESTERONE III | RIAP L3 | 2181EC | 98.3 | 9.75 | -0.001 | 135 | 0 | 98.292 | 4.223 | 4.296 | In use
C6000-A:E2-2 | PROGESTERONE III | RIAP L3 | 2147 | 101 | 10.1 | -0.585 | 148 | -0.058 | 95.094 | 3.469 | 3.648 | Study
C6000-A:E2-2 | PROGESTERONE III | RIAP L3 | 2110EC | 91.6 | 9.175 | -0.265 | 136 | -0.027 | 89.168 | 2.519 | 2.824 | Study
C6000-A:E2-2 | PROGESTERONE III | RIAP L1 | 2179EC | 3.49 | 0.35 | -0.379 | 123 | -0.038 | 3.357 | 0.387 | 11.538 | In use
C6000-A:E1-2 | PROLACTIN | RIAP L3 | 2110EC | 49.914 | 4.982 | 0.221 | 137 | 0.022 | 51.014 | 1.609 | 3.153 | Study
C6000-A:E1-2 | PROLACTIN | RIAP L3 | 2181EC | 73.132 | 7.3 | -0.539 | 133 | -0.054 | 69.197 | 3.95 | 5.708 | In use
C6000-A:E1-2 | PROLACTIN | RIAP L3 | 2147 | 46.7 | 4.7 | -0.379 | 153 | -0.038 | 44.917 | 2.273 | 5.06 | Study
C6000-A:E1-2 | PROLACTIN | RIAP L2 | 2107EC | 29.046 | 2.914 | -0.172 | 279 | -0.017 | 28.545 | 1.343 | 4.706 | Study
C6000-A:E1-2 | PROLACTIN | RIAP L2 | 2124EC | 29.09 | 2.182 | -1.034 | 78 | -0.078 | 26.833 | 2.22 | 8.272 | Study
C6000-A:E1-2 | PROLACTIN | RIAP L2 | 2228EC | 33.041 | 3.3 | -0.265 | 64 | -0.026 | 32.167 | 0.539 | 1.674 | In use
C6000-A:E2-2 | PROSTATE-SPECIFIC AG | RIAP L2 | 2124EC | 12.6 | 1.9 | -0.051 | 99 | -0.008 | 12.504 | 0.469 | 3.75 | Study
C6000-A:E2-2 | PROSTATE-SPECIFIC AG | RIAP L2 | 2107EC | 9.68 | 1.21 | 0.123 | 340 | 0.015 | 9.828 | 0.728 | 7.41 | Study
C6000-A:E2-2 | PROSTATE-SPECIFIC AG | RIAP L2 | 2228EC | 12.2 | 1.525 | -0.049 | 77 | -0.006 | 12.125 | 0.355 | 2.93 | In use
C6000-A:E2-2 | PROSTATE-SPECIFIC AG | RIAP L1 | 2105EC | 1.6 | 0.269 | -0.058 | 350 | -0.01 | 1.584 | 0.114 | 7.166 | Study
C6000-A:E2-2 | PROSTATE-SPECIFIC AG | RIAP L1 | 2179EC | 1.63 | 0.205 | -0.204 | 146 | -0.026 | 1.588 | 0.061 | 3.81 | In use
C6000-A:C1 | PROTEIN CSF | PNPUC TPC3 | 673292 | 0.218 | 0.022 | -0.17 | 41 | -0.017 | 0.214 | 0.007 | 3.396 | In use
C6000-A:C1 | PROTEIN CSF | PNPUC TPC3 | 602712 | 0.205 | 0.021 | 0.42 | 62 | 0.043 | 0.214 | 0.007 | 3.389 | Study
C6000-A:C1 | PROTEIN CSF | PPPUC TPC3 | 602710 | 1.51 | 0.15 | -0.146 | 165 | -0.014 | 1.488 | 0.021 | 1.389 | Study
C6000-A:C1 | PROTEIN CSF | PPPUC TPC3 | 531181 | 1.46 | 0.15 | -0.201 | 35 | -0.021 | 1.43 | 0.015 | 1.081 | Study
C6000-A:C1 | PROTEIN CSF | PNPUC TPC3 | 531179 | 0.21 | 0.021 | 0.057 | 79 | 0.006 | 0.211 | 0.006 | 2.808 | Study
C6000-A:C1 | PROTEIN CSF | PPPUC TPC3 | 646927 | 1.55 | 0.16 | -0.186 | 132 | -0.019 | 1.52 | 0.026 | 1.718 | In use
C6000-A:C1 | PROTEIN CSF | PNPUC TPC3 | 634291 | 0.202 | 0.02 | 0.523 | 30 | 0.052 | 0.212 | 0.023 | 11.011 | Study
C6000-A:C1 | PROTEIN CSF | PNPUC TPC3 | 649811 | 0.202 | 0.02 | 0.391 | 186 | 0.039 | 0.21 | 0.008 | 3.89 | Study
C6000-A:C1 | PROTEIN CSF | PPPUC TPC3 | 637091 | 1.64 | 0.16 | -0.148 | 122 | -0.014 | 1.616 | 0.048 | 2.991 | Study
C6000-A:C1 | PROTEIN URINE | PP PUC | 64692700 | 1.29 | 0.103 | -0.257 | 92 | -0.021 | 1.264 | 0.024 | 1.92 | Study
C6000-A:C1 | PROTEIN URINE | PP PUC | 637091 | 1.37 | 0.11 | -0.256 | 122 | -0.021 | 1.342 | 0.027 | 2.016 | Study
C6000-A:C1 | PROTEIN URINE | PP PUC | 602710 | 1.26 | 0.1 | -0.213 | 162 | -0.017 | 1.239 | 0.025 | 2.046 | Study
C6000-A:C1 | PROTEIN URINE | PN PUC | 602712 | 0.171 | 0.014 | 0.485 | 120 | 0.04 | 0.178 | 0.006 | 3.139 | Study
C6000-A:C1 | PROTEIN URINE | Ran-Liq 3 | 1238UC | 0.71 | 0.07 | 0.419 | 51 | 0.041 | 0.739 | 0.029 | 3.91 | In use
C6000-A:C1 | PROTEIN URINE | Ran-Liq 2 | 1237UC | 0.13 | 0.01 | 0.54 | 50 | 0.042 | 0.135 | 0.006 | 4.105 | In use
C6000-A:C1 | PROTEIN URINE | Ran-Liq 2 | 1209UC | 0.14 | 0.014 | -0.219 | 66 | -0.022 | 0.137 | 0.01 | 7.484 | Study
C6000-A:C1 | PROTEIN URINE | PN PUC | 531179 | 0.175 | 0.014 | -0.027 | 81 | -0.002 | 0.175 | 0.005 | 3.148 | Study
C6000-A:C1 | PROTEIN URINE | Ran-Liq 3 | 1210UC | 0.764 | 0.077 | 0.069 | 66 | 0.007 | 0.769 | 0.024 | 3.056 | Study
C6000-A:C1 | PROTEIN URINE | PN PUC | 649811 | 0.169 | 0.014 | 0.271 | 203 | 0.022 | 0.173 | 0.01 | 5.783 | Study
C6000-A:C1 | PROTEIN URINE | PP PUC | 531181 | 1.31 | 0.1 | -0.584 | 38 | -0.045 | 1.252 | 0.012 | 0.951 | Study
LIAISON XL | Renin | RENIN L2 | 7530320 | 102.3 | 8.5 | 1.045 | 42 | 0.087 | 111.185 | 6.292 | 5.659 | Study
LIAISON XL | Renin | RENIN L1 | 7520320 | 22.4 | 1.867 | 1.378 | 49 | 0.115 | 24.972 | 1.392 | 5.574 | Study
C6000-B:C2 | RHD3 | PPA1c H | 563451 | 10.7 | 0.6 | -0.147 | 515 | -0.008 | 10.612 | 0.327 | 3.079 | Study
C6000-B:C2 | RHD3 | PPA1c H | 509113 | 10.8 | 0.32 | -0.23 | 51 | -0.007 | 10.726 | 0.238 | 2.216 | Study
C6000-B:C2 | RHD3 | PNA1c H | 509111 | 5.65 | 0.17 | 0.438 | 51 | 0.013 | 5.725 | 0.114 | 1.995 | Study
C6000-B:C2 | RHD3 | PPA1c H | 617429 | 9.82 | 0.59 | -0.638 | 222 | -0.038 | 9.444 | 0.258 | 2.733 | In use
C6000-B:C2 | RHD3 | PNA1c H | 617428 | 5.7 | 0.34 | -0.046 | 224 | -0.003 | 5.684 | 0.127 | 2.241 | In use
C6000-B:C2 | RHD3 | PNA1c H | 548925 | 5.68 | 0.34 | 0.019 | 344 | 0.001 | 5.687 | 0.144 | 2.539 | Study
C6000-B:C2 | RHD3 | PNA1c H | 563450 | 5.63 | 0.17 | 0.735 | 165 | 0.022 | 5.755 | 0.117 | 2.025 | Study
C6000-B:C2 | RHEUMATOID FACTOR | RFCO2 | 590051 | 50.9 | 2.5 | 0.592 | 116 | 0.029 | 52.38 | 1.65 | 3.15 | Study
C6000-B:C2 | RHEUMATOID FACTOR | RFCO1 | 661133 | 22.7 | 1.5 | 0.469 | 131 | 0.031 | 23.404 | 0.961 | 4.107 | In use
C6000-B:C2 | RHEUMATOID FACTOR | RFCO2 | 661134 | 51.8 | 2.6 | 0.575 | 134 | 0.029 | 53.296 | 1.986 | 3.727 | In use
C6000-B:C2 | RHEUMATOID FACTOR | RFCO1 | 590050 | 21.3 | 1.4 | 0.592 | 119 | 0.039 | 22.129 | 1.106 | 4.996 | Study
C6000-B:C2 | RHEUMATOID FACTOR | RFCO1 | 542412 | 20.2 | 1.3 | 0.888 | 33 | 0.057 | 21.355 | 0.493 | 2.306 | Study
C6000-B:C2 | RHEUMATOID FACTOR | RFCO2 | 542413 | 49.3 | 2.5 | 0.544 | 33 | 0.028 | 50.661 | 1.343 | 2.651 | Study
C6000-A:E2-2 | SERUM FOLATE III | RIAP L2 | 2228EC | 12.027 | 0.715 | 0.604 | 88 | 0.036 | 12.459 | 3.896 | 31.268 | In use
C6000-A:E2-2 | SERUM FOLATE III | RIAP L3 | 2147 | 22.6 | 2.7 | 0.203 | 198 | 0.024 | 23.148 | 1.61 | 6.955 | Study
C6000-A:E2-2 | SERUM FOLATE III | RIAP L3 | 2110EC | 22.9 | 2.75 | 0.163 | 159 | 0.02 | 23.348 | 1.397 | 5.985 | Study
C6000-A:E2-2 | SERUM FOLATE III | RIAP L3 | 2181EC | 21.4 | 2.55 | 0.578 | 174 | 0.069 | 22.875 | 3.401 | 14.869 | In use
C6000-A:E2-2 | SERUM FOLATE III | RIAP L2 | 2124EC | 11.9 | 1.004 | 0.606 | 98 | 0.051 | 12.508 | 1.755 | 14.032 | Study
C6000-A:E2-2 | SERUM FOLATE III | RIAP L2 | 2107EC | 12.027 | 0.715 | 0.15 | 352 | 0.009 | 12.134 | 0.967 | 7.969 | Study
C6000-A:E2-2 | SEX HORMONE BINDING GLOBULIN | RIAP L1 | 2179EC | 9.7 | 0.97 | 0.245 | 115 | 0.025 | 9.938 | 0.336 | 3.383 | In use
C6000-A:E2-2 | SEX HORMONE BINDING GLOBULIN | RIAP L1 | 2105EC | 10.93 | 0.58 | 0.159 | 297 | 0.008 | 11.022 | 0.551 | 5 | Study
C6000-A:E2-2 | SEX HORMONE BINDING GLOBULIN | RIAP L2 | 2124EC | 39.3 | 4 | -0.247 | 82 | -0.025 | 38.312 | 1.807 | 4.717 | Study
C6000-A:E2-2 | SEX HORMONE BINDING GLOBULIN | RIAP L2 | 2107EC | 39 | 3.9 | -0.315 | 268 | -0.032 | 37.77 | 1.431 | 3.788 | Study
C6000-A:E2-2 | SEX HORMONE BINDING GLOBULIN | RIAP L2 | 2228EC | 34.1 | 3.4 | 0.418 | 66 | 0.042 | 35.521 | 0.879 | 2.476 | In use
C6000-A:ISE1 | SODIUM | s-PCCC 2 | 535719 | 135 | 4 | 0.843 | 193 | 0.025 | 138.373 | 1.892 | 1.367 | In use
C6000-B:ISE1 | SODIUM | s-PCCC 1-B | 525205 | 112 | 3 | 0.409 | 302 | 0.011 | 113.228 | 2.15 | 1.899 | In use
C6000-B:ISE1 | SODIUM | s-PCCC 2-B | 535719 | 135 | 4 | 0.572 | 188 | 0.017 | 137.287 | 5.765 | 4.199 | In use
C6000-A:ISE1 | SODIUM | s-PCCC 1 | 525027 | 115 | 2.1 | -0.273 | 588 | -0.005 | 114.427 | 1.847 | 1.614 | Study
C6000-A:ISE1 | SODIUM | PCCC 2 | 461597 | 134 | 4 | -0.116 | 88 | -0.003 | 133.534 | 1.819 | 1.362 | Study
C6000-A:ISE1 | SODIUM | s-PCCC 1 | 525205 | 112 | 3 | 0.745 | 342 | 0.02 | 114.234 | 2.005 | 1.755 | In use
C6000-B:ISE1 | SODIUM | PCCC 2-B | 519198 | 136 | 4 | 0.217 | 279 | 0.006 | 136.867 | 1.856 | 1.356 | Study
C6000-B:ISE1 | SODIUM | PCCC 1-B | 525027 | 115 | 2.1 | -0.292 | 369 | -0.005 | 114.388 | 1.74 | 1.521 | Study
C6000-B:ISE1 | SODIUM | PCCC 2-B | 461597 | 134 | 4 | -0.008 | 94 | 0 | 133.968 | 1.492 | 1.113 | Study
C6000-A:ISE1 | SODIUM | PCCC 1 | 525027 | 115 | 2.1 | -0.526 | 106 | -0.01 | 113.896 | 1.788 | 1.57 | Study
C6000-B:ISE1 | SODIUM | s-PCCC 2-B | 519198 | 136 | 4 | 0.233 | 697 | 0.007 | 136.933 | 2.021 | 1.476 | Study
C6000-B:ISE1 | SODIUM | s-PCCC 1-B | 525027 | 115 | 2.1 | -0.277 | 560 | -0.005 | 114.418 | 1.791 | 1.565 | Study
C6000-A:ISE1 | SODIUM | s-PCCC 2 | 519198 | 136 | 4 | 0.345 | 714 | 0.01 | 137.38 | 1.955 | 1.423 | Study
C6000-B:ISE1 | SODIUM URINE | Ran-Liq 2 | 1209UC | 73.6 | 4.4 | -0.664 | 241 | -0.04 | 70.676 | 9.898 | 14.005 | Study
C6000-B:ISE1 | SODIUM URINE | Ran-Liq 3 | 1210UC | 167 | 10 | -0.167 | 252 | -0.01 | 165.329 | 12.734 | 7.702 | Study
C6000-B:ISE1 | SODIUM URINE | Bio-liq 2 | 88162 | 165 | 9 | 0.259 | 164 | 0.014 | 167.335 | 3.526 | 2.107 | Study
C6000-B:ISE1 | SODIUM URINE | Ran-Liq 3 | 1238UC | 173 | 10.5 | -0.044 | 155 | -0.003 | 172.542 | 2.917 | 1.691 | In use
C6000-B:ISE1 | SODIUM URINE | Ran-Liq 2 | 1237UC | 86.8 | 5.2 | -0.361 | 138 | -0.022 | 84.92 | 2.277 | 2.681 | In use
C6000-B:ISE1 | SODIUM URINE | Bio-liq1 | 88161 | 77.4 | 4.85 | 0.035 | 165 | 0.002 | 77.57 | 1.961 | 2.528 | Study
C6000-A:E2-1 | TESTOSTERONE | RIAP L3 | 2181EC | 28.7 | 2.85 | -0.425 | 129 | -0.042 | 27.49 | 1.024 | 3.726 | In use
C6000-A:E2-1 | TESTOSTERONE | RIAP L2 | 2228EC | 14.7 | 1.45 | 0.171 | 63 | 0.017 | 14.948 | 0.476 | 3.181 | In use
C6000-A:E2-1 | TESTOSTERONE | RIAP L3 | 2147 | 28.7 | 2.85 | -0.369 | 145 | -0.037 | 27.647 | 1.631 | 5.899 | Study
C6000-A:E2-1 | TESTOSTERONE | RIAP L2 | 2107EC | 12.9 | 1.3 | -0.014 | 273 | -0.001 | 12.882 | 0.938 | 7.282 | Study
C6000-A:E2-1 | TESTOSTERONE | RIAP L2 | 2124EC | 14.6 | 1.45 | -0.153 | 78 | -0.015 | 14.379 | 0.458 | 3.188 | Study
C6000-A:E2-1 | TESTOSTERONE | RIAP L3 | 2110EC | 28.4 | 2.85 | 0.188 | 134 | 0.019 | 28.937 | 2.009 | 6.944 | Study
C6000-A:E1-1 | THYROGLOBULIN v2 | RIAP L3 | 2147 | 66.3 | 8.3 | -0.088 | 153 | -0.011 | 65.571 | 2.823 | 4.306 | Study
C6000-A:E1-1 | THYROGLOBULIN v2 | RIAP L1 | 2105EC | 2.15 | 0.12 | -0.424 | 302 | -0.024 | 2.099 | 0.108 | 5.159 | Study
C6000-A:E1-1 | THYROGLOBULIN v2 | RIAP L3 | 2181EC | 60 | 7.5 | 0.276 | 130 | 0.035 | 62.073 | 6.418 | 10.34 | In use
C6000-A:E1-1 | THYROGLOBULIN v2 | RIAP L3 | 2110EC | 68.8 | 8.6 | -0.207 | 144 | -0.026 | 67.02 | 3.168 | 4.727 | Study
C6000-A:E1-1 | THYROGLOBULIN v2 | RIAP L1 | 2179EC | 1.89 | 0.235 | 0.208 | 134 | 0.026 | 1.939 | 0.214 | 11.056 | In use
C6000-B:E3-1 | THYROID STIMULATING HORMONE | RIAP L2-B | 2124EC | 2.86 | 0.285 | 0.368 | 90 | 0.037 | 2.965 | 0.22 | 7.416 | Study
C6000-B:E3-1 | THYROID STIMULATING HORMONE | RIAP L1-B | 2179EC | 0.136 | 0.014 | 0.002 | 141 | 0 | 0.136 | 0.008 | 6.233 | In use
C6000-A:E2-2 | THYROID STIMULATING HORMONE | RIAP L2 | 2228EC | 2.54 | 0.255 | 0.214 | 98 | 0.022 | 2.595 | 0.067 | 2.576 | In use
C6000-B:E3-1 | THYROID STIMULATING HORMONE | RIAP L1-B | 2105EC | 0.104 | 0.006 | 0.4 | 316 | 0.023 | 0.106 | 0.005 | 5.018 | Study
C6000-A:E2-2 | THYROID STIMULATING HORMONE | RIAP L2 | 2107EC | 2.34 | 0.235 | 0.25 | 420 | 0.025 | 2.399 | 0.144 | 6.018 | Study
C6000-A:E2-2 | THYROID STIMULATING HORMONE | RIAP L2 | 2124EC | 2.86 | 0.17 | 0.413 | 119 | 0.025 | 2.93 | 0.142 | 4.83 | Study
C6000-B:E3-1 | THYROID STIMULATING HORMONE | RIAP L2-B | 2107EC | 2.34 | 0.235 | 0.222 | 312 | 0.022 | 2.392 | 0.141 | 5.914 | Study
C6000-B:E3-1 | THYROID STIMULATING HORMONE | RIAP L2-B | 2228EC | 2.54 | 0.255 | 0.199 | 77 | 0.02 | 2.591 | 0.111 | 4.283 | In use
C6000-A:E2-2 | THYROID STIMULATING HORMONE | RIAP L1 | 2179EC | 0.136 | 0.014 | -0.201 | 174 | -0.02 | 0.133 | 0.009 | 6.438 | In use
C6000-A:E2-2 | THYROID STIMULATING HORMONE | RIAP L1 | 2105EC | 0.104 | 0.006 | -0.102 | 436 | -0.006 | 0.103 | 0.005 | 4.785 | Study
C6000-B:C2 | TOTAL BILIRUBIN | PCCC 2-B | 519198 | 61.8 | 3.7 | -0.478 | 397 | -0.029 | 60.03 | 2.908 | 4.845 | Study
C6000-A:C1 | TOTAL BILIRUBIN | MAS BILIRUBIN 3 | BC23063A | 327 | 17.8 | -0.318 | 73 | -0.017 | 321.348 | 11.243 | 3.499 | Study
C6000-A:C1 | TOTAL BILIRUBIN | PCCC 2 | 519198 | 61.8 | 3.7 | -0.282 | 417 | -0.017 | 60.758 | 2.322 | 3.822 | Study
C6000-B:C2 | TOTAL BILIRUBIN | PCCC 2-B | 535719 | 62.3 | 3.7 | -0.261 | 99 | -0.015 | 61.335 | 1.826 | 2.977 | In use
C6000-A:C1 | TOTAL BILIRUBIN | PCCC 2 | 535719 | 62.3 | 3.7 | 0.264 | 108 | 0.016 | 63.277 | 2.494 | 3.941 | In use
C6000-B:C2 | TOTAL BILIRUBIN | PCCC 1-B | 525205 | 16.4 | 1 | 0.027 | 166 | 0.002 | 16.427 | 0.764 | 4.649 | In use
C6000-A:C1 | TOTAL BILIRUBIN | PCCC 2 | 461597 | 63.5 | 3.8 | -0.425 | 77 | -0.025 | 61.884 | 1.907 | 3.082 | Study
C6000-A:C1 | TOTAL BILIRUBIN | PCCC 1 | 525205 | 16.4 | 1 | 0.061 | 168 | 0.004 | 16.461 | 0.741 | 4.504 | In use
C6000-A:C1 | TOTAL BILIRUBIN | PCCC 1 | 525027 | 16.2 | 1 | -0.429 | 451 | -0.026 | 15.771 | 0.882 | 5.593 | Study
C6000-B:C2 | TOTAL BILIRUBIN | PCCC 2-B | 461597 | 63.5 | 3.8 | -0.232 | 78 | -0.014 | 62.617 | 1.121 | 1.79 | Study
C6000-A:C1 | TOTAL BILIRUBIN | MAS BILIRUBIN 3 | BC24033A | 340 | 22.7 | 0.273 | 414 | 0.018 | 346.205 | 9.544 | 2.757 | In use
C6000-B:C2 | TOTAL BILIRUBIN | MAS BILIRUBIN-B | BC23063A | 327 | 17.8 | -0.16 | 72 | -0.009 | 324.158 | 9.218 | 2.844 | Study
C6000-B:C2 | TOTAL BILIRUBIN | MAS BILIRUBIN-B | BC24033A | 340 | 22.7 | 0.102 | 385 | 0.007 | 342.326 | 11.669 | 3.409 | In use
C6000-B:C2 | TOTAL BILIRUBIN | PCCC 1-B | 525027 | 16.2 | 1 | -0.54 | 418 | -0.033 | 15.66 | 0.751 | 4.798 | Study
C6000-A:C1 | TOTAL CHOLESTEROL | PCCC 1 | 525027 | 2.43 | 0.12 | -0.434 | 93 | -0.021 | 2.378 | 0.056 | 2.353 | Study
C6000-A:C1 | TOTAL CHOLESTEROL | s-PCCC 2 | 519198 | 4.48 | 0.22 | -0.061 | 610 | -0.003 | 4.467 | 0.087 | 1.944 | Study
C6000-B:C2 | TOTAL CHOLESTEROL | s-PCCC 1-B | 525205 | 2.42 | 0.12 | 0.027 | 239 | 0.001 | 2.423 | 0.058 | 2.401 | In use
C6000-B:C2 | TOTAL CHOLESTEROL | s-PCCC 2-B | 519198 | 4.48 | 0.22 | 0.044 | 604 | 0.002 | 4.49 | 0.088 | 1.966 | Study
C6000-B:C2 | TOTAL CHOLESTEROL | PCCC 2-B | 461597 | 4.48 | 0.22 | 0.263 | 64 | 0.013 | 4.538 | 0.059 | 1.289 | Study
C6000-A:C1 | TOTAL CHOLESTEROL | s-PCCC 1 | 525205 | 2.42 | 0.12 | 0.061 | 264 | 0.003 | 2.427 | 0.054 | 2.244 | In use
C6000-A:C1 | TOTAL CHOLESTEROL | s-PCCC 2 | 535719 | 4.4 | 0.22 | -0.241 | 165 | -0.012 | 4.347 | 0.102 | 2.356 | In use
C6000-B:C2 | TOTAL CHOLESTEROL | PCCC 1-B | 525027 | 2.43 | 0.12 | -0.259 | 76 | -0.013 | 2.399 | 0.04 | 1.686 | Study
C6000-A:C1 | TOTAL CHOLESTEROL | s-PCCC 1 | 525027 | 2.43 | 0.12 | -0.049 | 521 | -0.002 | 2.424 | 0.059 | 2.454 | Study
C6000-B:C2 | TOTAL CHOLESTEROL | s-PCCC 2-B | 535719 | 4.4 | 0.22 | -0.407 | 158 | -0.02 | 4.31 | 0.072 | 1.663 | In use
C6000-B:C2 | TOTAL CHOLESTEROL | s-PCCC 1-B | 525027 | 2.43 | 0.12 | 0.02 | 505 | 0.001 | 2.432 | 0.048 | 1.97 | Study
C6000-A:C1 | TOTAL CHOLESTEROL | PCCC 2 | 461597 | 4.55 | 0.23 | -0.281 | 76 | -0.014 | 4.485 | 0.098 | 2.195 | Study
C6000-A:C1 | TOTAL PROTEIN | PCCC 1 | 525027 | 48.1 | 1.9 | -0.023 | 408 | -0.001 | 48.055 | 0.776 | 1.616 | Study
C6000-A:C1 | TOTAL PROTEIN | PCCC 2 | 461597 | 81.1 | 2.2 | -0.1 | 73 | -0.003 | 80.881 | 1.084 | 1.34 | Study
C6000-B:C2 | TOTAL PROTEIN | s-PCCC 1-B | 525027 | 48.1 | 1.9 | -0.213 | 42 | -0.008 | 47.695 | 0.534 | 1.12 | Study
C6000-B:C2 | TOTAL PROTEIN | PCCC 1-B | 525205 | 48.2 | 1.9 | -0.49 | 151 | -0.019 | 47.27 | 1.092 | 2.31 | In use
C6000-A:C1 | TOTAL PROTEIN | PCCC 2 | 519198 | 77.8 | 3.1 | 0.122 | 387 | 0.005 | 78.18 | 1.238 | 1.584 | Study
C6000-B:C2 | TOTAL PROTEIN | PCCC 2-B | 519198 | 77.8 | 3.1 | -0.164 | 327 | -0.007 | 77.293 | 1.561 | 2.019 | Study
C6000-B:C2 | TOTAL PROTEIN | s-PCCC 2-B | 461597 | 81.1 | 2.2 | -0.408 | 38 | -0.011 | 80.203 | 0.872 | 1.087 | Study
C6000-B:C2 | TOTAL PROTEIN | PCCC 2-B | 535719 | 76.6 | 3.1 | -0.79 | 89 | -0.032 | 74.152 | 1.754 | 2.365 | In use
C6000-A:C1 | TOTAL PROTEIN | PCCC 2 | 535719 | 76.6 | 3.1 | -0.345 | 99 | -0.014 | 75.53 | 2.334 | 3.091 | In use
C6000-B:C2 | TOTAL PROTEIN | PCCC 1-B | 525027 | 48.1 | 1.9 | -0.272 | 277 | -0.011 | 47.584 | 0.915 | 1.924 | Study
C6000-A:C1 | TOTAL PROTEIN | PCCC 1 | 525205 | 48.2 | 1.9 | 0.03 | 164 | 0.001 | 48.256 | 1.371 | 2.842 | In use
C6000-A:C1 | TRANSFERRIN | s-PCCC 1 | 525027 | 1.98 | 0.12 | -0.218 | 54 | -0.013 | 1.954 | 0.065 | 3.332 | Study
C6000-A:C1 | TRANSFERRIN | s-PCCC 2 | 461597 | 3.49 | 0.21 | -0.465 | 48 | -0.028 | 3.392 | 0.135 | 3.989 | Study
C6000-A:C1 | TRANSFERRIN | PCCC 2 | 535719 | 3.33 | 0.2 | 0.132 | 110 | 0.008 | 3.356 | 0.06 | 1.794 | In use
C6000-A:C1 | TRANSFERRIN | PCCC 2 | 519198 | 3.42 | 0.21 | 0.122 | 425 | 0.007 | 3.446 | 0.113 | 3.288 | Study
C6000-A:C1 | TRANSFERRIN | PCCC 1 | 525205 | 1.89 | 0.11 | 0.266 | 188 | 0.016 | 1.919 | 0.036 | 1.867 | In use
C6000-A:C1 | TRANSFERRIN | PCCC 1 | 525027 | 1.98 | 0.12 | 0.282 | 365 | 0.017 | 2.014 | 0.061 | 3.04 | Study
C6000-A:E1-2 | TRI-IODO THYRONINE | RIAP L2 | 2228EC | 15.6 | 1.95 | 0.021 | 65 | 0.003 | 15.641 | 0.397 | 2.536 | In use
C6000-A:E1-2 | TRI-IODO THYRONINE | RIAP L2 | 2107EC | 13.6 | 1.7 | 0.123 | 276 | 0.015 | 13.81 | 0.479 | 3.468 | Study
C6000-A:E1-2 | TRI-IODO THYRONINE | RIAP L2 | 2124EC | 15.3 | 0.86 | -0.159 | 74 | -0.009 | 15.164 | 0.541 | 3.567 | Study
C6000-A:E1-2 | TRI-IODO THYRONINE | RIAP L1 | 2179EC | 4.96 | 0.62 | 0.148 | 117 | 0.018 | 5.052 | 0.246 | 4.878 | In use
C6000-A:E1-2 | TRI-IODO THYRONINE | RIAP L1 | 2105EC | 4.4 | 0.249 | 0.011 | 294 | 0.001 | 4.403 | 0.245 | 5.556 | Study
C6000-A:C1 | TRIGLYCERIDE | PCCC 1 | 525027 | 1.32 | 0.07 | 0.436 | 341 | 0.023 | 1.35 | 0.047 | 3.507 | Study
C6000-A:C1 | TRIGLYCERIDE | PCCC 2 | 519198 | 2.43 | 0.12 | 0.193 | 385 | 0.01 | 2.453 | 0.047 | 1.935 | Study
C6000-A:C1 | TRIGLYCERIDE | PCCC 2 | 535719 | 2.46 | 0.12 | -0.132 | 107 | -0.006 | 2.444 | 0.052 | 2.139 | In use
C6000-A:C1 | TRIGLYCERIDE | s-PCCC 1 | 525027 | 1.32 | 0.07 | 0.297 | 52 | 0.016 | 1.341 | 0.03 | 2.249 | Study
C6000-A:C1 | TRIGLYCERIDE | s-PCCC 2 | 461597 | 2.43 | 0.12 | 0.109 | 42 | 0.005 | 2.443 | 0.046 | 1.869 | Study
C6000-A:C1 | TRIGLYCERIDE | PCCC 1 | 525205 | 1.34 | 0.07 | 0.046 | 166 | 0.002 | 1.343 | 0.029 | 2.178 | In use
C6000-B:C2 | UREA | PCCC 2-B | 461597 | 20.2 | 1 | -0.523 | 77 | -0.026 | 19.677 | 0.37 | 1.878 | Study
C6000-A:C1 | UREA | s-PCCC 1 | 525027 | 6.66 | 0.33 | 0.146 | 546 | 0.007 | 6.708 | 0.178 | 2.652 | Study
C6000-B:C2 | UREA | PCCC 1-B | 525027 | 6.66 | 0.33 | -0.302 | 91 | -0.015 | 6.56 | 0.128 | 1.953 | Study
C6000-A:C1 | UREA | PCCC 2 | 461597 | 20.2 | 1 | -0.495 | 76 | -0.024 | 19.705 | 0.42 | 2.134 | Study
C6000-B:C2 | UREA | s-PCCC 1-B | 525027 | 6.66 | 0.33 | 0.357 | 530 | 0.018 | 6.778 | 0.181 | 2.666 | Study
C6000-A:C1 | UREA | s-PCCC 2 | 519198 | 20.7 | 1 | -0.294 | 644 | -0.014 | 20.406 | 0.54 | 2.645 | Study
C6000-B:C2 | UREA | s-PCCC 1-B | 525205 | 6.76 | 0.34 | -0.062 | 256 | -0.003 | 6.739 | 0.201 | 2.981 | In use
C6000-B:C2 | UREA | s-PCCC 2-B | 519198 | 20.7 | 1 | -0.029 | 639 | -0.001 | 20.671 | 0.517 | 2.502 | Study
C6000-A:C1 | UREA | s-PCCC 2 | 535719 | 20.6 | 1 | -0.651 | 167 | -0.032 | 19.949 | 0.626 | 3.139 | In use
C6000-A:C1 | UREA | s-PCCC 1 | 525205 | 6.76 | 0.34 | -0.469 | 267 | -0.024 | 6.6 | 0.208 | 3.154 | In use
C6000-B:C2 | UREA | s-PCCC 2-B | 535719 | 20.6 | 1 | -0.549 | 164 | -0.027 | 20.051 | 0.436 | 2.175 | In use
C6000-A:C1 | UREA | PCCC 1 | 525027 | 6.66 | 0.33 | -0.073 | 95 | -0.004 | 6.636 | 0.127 | 1.915 | Study
C6000-B:C2 | UREA URINE | Ran-Liq 2 | 1209UC | 145 | 14.5 | 0.325 | 194 | 0.032 | 149.709 | 5.062 | 3.381 | Study
C6000-B:C2 | UREA URINE | Bio-liq1 | 88161 | 149 | 9.5 | 0.264 | 136 | 0.017 | 151.508 | 5.226 | 3.45 | Study
C6000-B:C2 | UREA URINE | Ran-Liq 3 | 1210UC | 274 | 27.5 | 0.151 | 229 | 0.015 | 278.159 | 15.052 | 5.411 | Study
C6000-B:C2 | UREA URINE | Bio-liq 2 | 88162 | 253 | 16 | 0.29 | 131 | 0.018 | 257.644 | 9.902 | 3.843 | Study
C6000-A:C1 | UREA URINE | Ran-Liq 2 | 1237UC | 149 | 15 | 0.024 | 50 | 0.002 | 149.36 | 5.772 | 3.865 | In use
C6000-A:C1 | UREA URINE | Ran-Liq 3 | 1238UC | 289 | 29 | -0.258 | 49 | -0.026 | 281.529 | 12.017 | 4.269 | In use
C6000-B:C2 | UREA URINE | Ran-Liq 2 | 1237UC | 149 | 15 | 0.039 | 122 | 0.004 | 149.59 | 5.571 | 3.724 | In use
C6000-A:C1 | UREA URINE | Ran-Liq 3 | 1210UC | 274 | 27.5 | 0.049 | 73 | 0.005 | 275.344 | 21.418 | 7.779 | Study
C6000-A:C1 | UREA URINE | Ran-Liq 2 | 1209UC | 145 | 14.5 | 0.588 | 73 | 0.059 | 153.532 | 25.355 | 16.514 | Study
C6000-B:C2 | UREA URINE | Ran-Liq 3 | 1238UC | 289 | 29 | -0.124 | 137 | -0.012 | 285.408 | 9.087 | 3.184 | In use
C6000-A:C1 | URIC ACID | PCCC 1 | 525205 | 0.282 | 0.014 | 0.342 | 141 | 0.017 | 0.287 | 0.006 | 2.129 | In use
C6000-A:C1 | URIC ACID | PCCC 1 | 525027 | 0.285 | 0.014 | 0.314 | 357 | 0.015 | 0.289 | 0.006 | 2.198 | Study
C6000-B:C2 | URIC ACID | PCCC 2-B | 535719 | 0.595 | 0.03 | 0.029 | 79 | 0.001 | 0.596 | 0.012 | 2.095 | In use
C6000-B:C2 | URIC ACID | PCCC 1-B | 525027 | 0.285 | 0.014 | 0.14 | 255 | 0.007 | 0.287 | 0.006 | 2.063 | Study
C6000-B:C2 | URIC ACID | s-PCCC 2-B | 461597 | 0.569 | 0.028 | 0.005 | 33 | 0 | 0.569 | 0.005 | 0.922 | Study
C6000-A:C1 | URIC ACID | PCCC 2 | 461597 | 0.569 | 0.028 | 0.054 | 67 | 0.003 | 0.571 | 0.009 | 1.574 | Study
C6000-A:C1 | URIC ACID | PCCC 2 | 519198 | 0.576 | 0.029 | 0.412 | 326 | 0.021 | 0.588 | 0.012 | 2.119 | Study
C6000-B:C2 | URIC ACID | PCCC 1-B | 525205 | 0.282 | 0.014 | 0.192 | 134 | 0.01 | 0.285 | 0.006 | 2.003 | In use
C6000-B:C2 | URIC ACID | PCCC 2-B | 519198 | 0.576 | 0.029 | 0.287 | 300 | 0.014 | 0.584 | 0.012 | 2.009 | Study
C6000-A:C1 | URIC ACID | PCCC 2 | 535719 | 0.595 | 0.03 | 0.349 | 93 | 0.018 | 0.605 | 0.012 | 2.011 | In use
C6000-B:C2 | URIC ACID | s-PCCC 1-B | 525027 | 0.285 | 0.014 | 0.083 | 38 | 0.004 | 0.286 | 0.003 | 1.104 | Study
C6000-B:C2 | URIC ACID URINE | Bio-liq1 | 88161 | 0.847 | 0.05 | 0.139 | 104 | 0.008 | 0.854 | 0.021 | 2.474 | Study
C6000-B:C2 | URIC ACID URINE | Bio-liq 2 | 88162 | 1.27 | 0.07 | 0.176 | 99 | 0.01 | 1.282 | 0.029 | 2.279 | Study
C6000-B:C2 | URIC ACID URINE | Ran-Liq 3 | 1238UC | 0.85 | 0.08 | 0.261 | 117 | 0.025 | 0.871 | 0.019 | 2.163 | In use
C6000-B:C2 | URIC ACID URINE | Ran-Liq 2 | 1237UC | 0.53 | 0.05 | 0.293 | 108 | 0.028 | 0.545 | 0.014 | 2.637 | In use
C6000-B:C2 | URIC ACID URINE | Ran-Liq 3 | 1210UC | 0.87 | 0.09 | -0.105 | 198 | -0.011 | 0.861 | 0.04 | 4.613 | Study
C6000-B:C2 | URIC ACID URINE | Ran-Liq 2 | 1209UC | 0.54 | 0.05 | 0.018 | 197 | 0.002 | 0.541 | 0.039 | 7.118 | Study
OSMO1 | Urine osmolality | Biorad Urine QC1 | 88162 | 789 | 5.667 | -0.22 | 93 | -0.002 | 787.753 | 3.236 | 0.411 | Study
OSMO1 | Urine osmolality | Ran-Liq 2 | 1237UC | 437.063 | 5.905 | -0.156 | 79 | -0.002 | 436.139 | 5.126 | 1.175 | In use
OSMO1 | Urine osmolality | Biorad Urine QC1 | 88161 | 437 | 3.5 | -0.143 | 102 | -0.001 | 436.5 | 1.913 | 0.438 | Study
OSMO1 | Urine osmolality | Randox Urine QC 3 | 1210UC | 804.8 | 4.57 | 0.507 | 34 | 0.003 | 807.118 | 4.721 | 0.585 | Study
OSMO1 | Urine osmolality | Ran-Liq 3 | 1238UC | 829.477 | 6.181 | -0.103 | 87 | -0.001 | 828.839 | 5.754 | 0.694 | In use
OSMO1 | Urine osmolality | Randox Urine QC 2 | 1209UC | 414 | 41.5 | 0.028 | 32 | 0.003 | 415.156 | 2.807 | 0.676 | Study
C6000-A:E1-2 | VITAMIN B12 II | RIAP L2 | 2228EC | 417 | 41.5 | 0.628 | 91 | 0.062 | 443.052 | 15.306 | 3.455 | In use
C6000-A:E1-2 | VITAMIN B12 II | RIAP L1 | 2105EC | 167 | 16.7 | 0.241 | 436 | 0.024 | 171.022 | 12.719 | 7.437 | Study
C6000-A:E1-2 | VITAMIN B12 II | RIAP L2 | 2107EC | 430 | 43 | 0.264 | 407 | 0.026 | 441.334 | 17.812 | 4.036 | Study
C6000-A:E1-2 | VITAMIN B12 II | RIAP L1 | 2179EC | 154 | 15.5 | 0.352 | 190 | 0.035 | 159.453 | 9.974 | 6.255 | In use
C6000-A:E1-2 | VITAMIN B12 II | RIAP L2 | 2124EC | 439 | 44 | -0.061 | 115 | -0.006 | 436.328 | 36.8 | 8.434 | Study